INTERNATIONAL ELECTROTECHNICAL COMMISSION SYSTEM FOR CERTIFICATION TO STANDARDS RELATING TO EQUIPMENT FOR

USE IN EXPLOSIVE ATMOSPHERES (IECEx SYSTEM)

Title: 	Draft Operational Document IECEx OD 024 Version 03/Toronto IECEx Rules of Procedure covering testing, or witnessing testing at a manufacturer’s or user’s facility

Circulated to: ExTAG Members

## INTRODUCTION

During the ExTAG Shanghai 2006 Meeting it was recommended that Ex/OD024/V1

Rules of Procedure for testing at other locations be withdrawn and that further work be undertaken by the ExTAG Working Group responsible, WG06.

During the ExTAG 2008 Paris meeting a new Convener of ExTAG WG6, Mr Nick Maalouf of QPS was appointed. At the time Mr Maalouf asked, and was given approval for, the following proposals:

- Terms of Reference for the ExTAG WG6 be revised as follows:

To establish the mechanisms by which test data generated by a 			manufacturer under the control of an ExTL or ExCB can be accepted 		within the IECEx Scheme can be accommodated

- A work program for WG6 be developed

This document is a revised version of OD024/V1 and has been prepared as a result of a ExTAG WG6 Meeting held on April 22/23, 2009, in Toronto, and subsequent work carried out by the WG. It is issued for consideration during the ExTAG Melbourne Meeting. For Annex A referenced in this document see ExTAG/178/CD

Michel Brenon

ExTAG Secretary

| Address: SAI Global Building 286 Sussex Street Sydney NSW 2000 Australia | IECEx Secretariat Contact Details: Tel: +61 2 8206 6940 Fax: +61 2 8206 6272 e-mail: chris.agius@iecex.com http://www.iecex.com | ExTAG Secretary Mr Michel Brenon/LCIE Tel: +33 1 40 95 5519 Fax: +33 1 40 95 5520 e-mail michel.brenon@lcie.fr |
| --- | --- | --- |

XXX  IEC:200X	– 2 –	Draft OD024 Version 03/Toronto

July 2009

IECEx OPERATIONAL DOCUMENT

Draft IECEx OD 024 Version 03/Toronto

Title: IECEx Rules of Procedure covering testing, or witnessing testing at a manufacturer’s or user’s facility

| Clause | Title | Page |
| --- | --- | --- |
| 1 | INTRODUCTION | 2 |
| 2 | DEFINITIONS | 2 |
| 3 | SCOPE | 3 |
| 4 | OVERALL REQUIREMENTS | 3 |
| 5 | REQUIREMENTS AND PROCEDURES FOR OFF-SITE TESTING | 3 |
| 5.1 | General Requirements | 3 |
| 5.2 | Role and Responsibilities of the ExTL | 4 |
| 5.3 | Role and Responsibilities of the ExCB | 4 |
| 5.4 | Role and Responsibilities of the Manufacturer or User Test Facility | 4 |
| 5.5 | Choice of ExCBs/ExTLs by the Manufacturer or User Test Facility | 5 |
| 5.6 | Site Assessment/Reassessment of the Manufacturer or User Test Facility | 5 |
| 5.7 | Requirements for Conducting Tests and Issuing Test Reports and Certificates | 6 |
| 6 | REQUIREMENTS AND PROCEDURES FOR WITNESS TESTING | 6 |
| 6.1 | General Requirements | 6 |
| 6.2 | Role and Responsibilities of the ExTL | 7 |
| 6.3 | Role and Responsibilities of the ExCB | 7 |
| 6.4 | Role and Responsibilities of the Manufacturer or User Test Facility | 8 |
| 6.5 | Choice of ExCBs/ExTLs by the Manufacturer or User Test Facility | 8 |
| 6.6 | Site Assessment/Reassessment of the Manufacturer or User Test Facility | 8 |
| 6.7 | Requirements for Conducting Tests and Issuing Test Reports and Certificates | 9 |
| 6.8 | Long Term Tests | 10 |
| 6.9 | Remote Witnessing of Testing | 10 |
| 6.10 | Partially Witnessed Testing | 10 |

# INTRODUCTION

## There are various situations that necessitate that testing be conducted at a facility other than an ExTL. Examples of such situations are:

- Some manufacturers and end users who use the IECEx Scheme have the capability in terms of personnel, facility and equipment to conduct a part of or all the tests stipulated in a given standard within the scope of the IECEx System.

- The size of the product intended for testing is too large to bring to, and handle at the ExTL premises.

- The product forms, and is to be tested as, part of an installation located at the end user site.

## In recognition of the above conditions, and in response to identified needs to utilize existing capabilities other than those of an ExTL, this operational document has been established for testing to be conducted at such facilities under controlled conditions as specifically detailed and defined in this Operational Document.

## Off-site testing and witness testing are currently being employed by ExTLs and ExCBs to accommodate the above situations. This Operational Document has been developed to recognize such existing practices and to provide a common approach so as to ensure consistency, transparency, and proper controls, and to maintain the credibility of the IECEx System

## It is the intent of this Operational Document that testing carried out under these programs be performed with the same rigor as testing conducted at an ExTL

## The activities of an ExTL defined in this Operational Document are carried out in conjunction with an ExCB, in accordance with the requirements of IECEx02.

# DEFINITIONS

For the purpose of this Operational Document the following definitions apply:

ExCB Staff - Staff of an ExCB identified in ExCB competency records as being qualified to review test and assessment reports for the Test Standard in question.

ExTL Staff - Staff of an ExTL identified in ExTL technical competency records as being qualified to conduct testing to the requirements of the Test Standard in question.

Long Term Tests - Tests of a duration exceeding 7 hours without intervention by ExTL testing staff between the start and the end of the test.

Manufacturer Test Facility - Test facility owned by a Manufacturer which is being used to test specified products over which the manufacturer has complete production control.

Off-Site Testing - Program/procedure under which testing is conducted by ExTL staff at a manufacturer’s or user’s facility.

Partially Witnessed Testing – Testing witnessed by ExTL or ExCB staff only on some parts of the agreed upon testing program.

Remote Witness Test - Testing witnessed by ExTL staff using an electronic medium such as a video camera.

Site Assessment - Assessment carried out by ExTL or ExCB staff to verify the ability and capability of a manufacturer or user test facility to participate in the off-site testing and witness testing programs.

Test Standard - IEC standard for a specific protection method used to test the product in question at the manufacturer or user test facility, and which is used in the IECEx System.

Witness Testing – Program/procedure under which testing is conducted by personnel of the manufacturer or user, using their facility, equipment and personnel; and witnessed by ExCB or ExTL staff.

# SCOPE

## This operational document sets out the requirements covering the following programs:

- “Off-Site Testing”.

- “Witness Testing”.

## The requirements of this Operational Document apply to equipment covered by the IECEx System.

# OVERALL REQUIREMENTS

The following criteria apply to both off-site and witness testing programs.

## All testing carried out under these programs shall be performed in accordance with the following:

- The requirements of this Operational Document.

- The requirements of the applicable IECEx Operational Documents.

- The relevant requirements of ISO/IEC 17025 as defined under each program.

## The ExCB and ExTL wishing to use these programs must have the relevant Test Standard(s) included in their IECEx scope of recognition.

## The decision by an ExTL to conduct testing under these programs must be approved in advance by the ExCB.

## Participating users must own the participating test facility.

## Participating manufacturers must own the participating test facility; and own, or be in full control over all the production means of the product being tested.

# REQUIREMENTS AND PROCEDURES FOR OFF-SITE TESTING

## General Requirements

### Under this Operational Document, the following principles shall apply:

- Personnel of the manufacturer or user may assist in the preparation for tests and the conduct of agreed upon tests.

- The ExTL remains fully responsible for the test results along with the preparation and content of the required test report.

### An “Off-Site Testing Agreement” shall be signed between the ExCB or the ExTL, and the manufacturer or user. The agreement shall cover the following, as a minimum:

- The scope of work (specific tests or test standards) to be carried out at the participating test facility.

- The facilities and services to be provided by the manufacturer or user.

- The right of initial assessment and re-assessment of the manufacturer or user by the ExTL or the ExCB.

- Use of data generated.

## Role and Responsibilities of the ExTL

The ExTL shall:

### Have the required quality procedures which adequately cover all aspects of off-site testing.

### Ensure that its staff involved in off-site testing are provided the necessary initial and on-going training on the requirements of this operational document.

### Carry out the required initial and on-going site assessments.

### Define the role and responsibility of personnel involved in off-site testing on behalf of the manufacturer or user.

### Where necessary, verify the competence of the manufacturer or user personnel to support the off-site testing program.

### Perform testing at the manufacture or user test facility.

### Prepare and sign the required test reports.

### Sign the agreement stipulated in Clause 5.1.2 above.

## Role and Responsibilities of the ExCB

The ExCB shall be responsible for the following activities:

    - Assessment of the quality procedures of the ExTL, applicable to the operation of off-site testing.

    - Ensuring that its staff involved in off-site testing are provided the necessary initial and on-going training on the requirements of this operational document.

    - Review and approval of all pertinent documents generated under this operational document by the ExTL.

    - Ensuring there is an appropriate agreement between the ExCB or the ExTL and the manufacturer/ user, as per Clause 5.1.2 above.

    - Issuing ExTR based on off-site testing.

## Role and Responsibilities of the Manufacturer or User Test facility

## Participating manufacturer or user test facilities are responsible for:

### Making a formal application to participate in the program

### Signing the required agreement with the ExTL or ExCB.

### Demonstrating that the facilities are in compliance with the relevant requirements of ISO/IEC 17025 and the requirements of this program. Relevant requirements of ISO/IEC 17025 shall include the following elements, as a minimum:

- Critical consumables

- Personnel

- Accommodations and environmental conditions

- Equipment

- Measurement traceability / calibration

- Identification of test items

### Designating competent personnel for the testing activities conducted at their facility.

### Ensuring that assigned personnel follow the instructions of the ExTL staff conducting off-site testing at the facility.

## Choice of ExCBs/ExTLs by the Manufacturer or User Test Facility

A participating test facility shall be permitted to have an agreement with more than one ExCB for a given test standard.

## Site Assessment/Reassessment of the Manufacturer or User Test Facility

### ExTL Staff shall carry out the initial site assessment of the participating test facility in accordance with the ExTL’s procedure for off-site testing.

### The continued validity of the initial assessment or reassessment shall be evaluated during each subsequent off-site testing project/visit. The evaluation shall consider the following factors:

- The scope of work defined in the agreement referenced in clause 5.1.2

- The frequency of visits

- The extent of organizational or equipment changes at the facility

- Other changes that affect the performance of the participating test facility relative to the conduct of off-site testing.

### Where testing is conducted using test equipment of the ExTL, the site assessment shall cover relevant requirements of ISO/IEC 17025. Relevant requirements of ISO/IEC 17025 shall include the following elements. As a minimum:

- Control of environmental conditions at the facility.

- Stability and capacity of power supplies.

- Competence of personnel at the participating test facility assigned to support off-site testing activities.

- Identification of test items

### Where testing is conducted using test equipment provided by the participating test facility, the site assessment shall cover the relevant requirements of ISO/IEC 17025. Relevant requirements of ISO/IEC 17025 shall include the following elements, as a minimum:

- Critical consumables

- Accommodations and environmental conditions.

- Stability and capacity of power supplies.

- Availability and adequacy of test equipment.

- Measurement traceability/calibration.

- Competence of personnel at the participating test facility assigned to support off-site testing activities.

- Identification of test items

### The ExTL shall document the results of any assessment/reassessment activity using the appropriate Assessment Report Form (Annex A), and provide a copy of the report to the ExCB for review and approval.

### All documentation related to assessment/reassessment activities shall be made available as follows:

- For audit by the ExCB.

- For review during any subsequent peer reassessments conducted under the IECEx System.

- To the IECEx Secretariat upon request.

## Requirements for Conducting Tests and Issuing Test Reports and Certificates

### All testing carried out under the off-site testing program shall be performed with the same rigor as testing conducted at the ExTL, and using the same test procedures used by the ExTL.

### Where testing is performed using any test equipment of the participating test facility, the ExTL staff shall:

- Verify that the test instrument is properly calibrated by an ISO 17025 accredited calibration test lab, has calibration that is traceable to National Standards,

- Verify that the test instrument is suitable for making the required measurement and has the needed accuracy as required by the test standard involved.

- Become familiar with the use and operation of the instrument.

- Ensure that the instrument is functioning.

### Where personnel of the participating test facility are assisting in testing activity (such as placement of thermocouples, recording test results, etc.) the ExTL staff shall:

- Provide clear and concise instructions as to the type and extent of assistance provided.

- Examine the work done prior to commencing and during testing.

### Each test sample shall be uniquely identified.

### Test reports/records prepared shall meet the following criteria:

- Be the relevant IECEx Test Report Forms.

- Be signed /reviewed in accordance with the relevant procedures of the IECEx System.

- Record the date of testing and identification of the all test samples.

- Indicate the name and address of the manufacturer or user test facility used

- Indicate what tests have been carried out at the manufacturer or user test facility.

- Record all test instruments used in testing.

- Incorporate applicable test data.

### When an ExTR is based on off-site testing, a reference to the use of the manufacturer or user test facility shall be included as “Additional Information” in the on-line ExTR summary.

# REQUIREMENTS AND PROCEDURES FOR WITNESS TESTING

## General Requirements

### Under this program, the following principles shall apply:

- The manufacturer or user test facility uses its own test equipment.

- Testing is carried out by personnel of the manufacturer or user test facility.

- The ExCB or the ExTL involved witnesses the testing.

- The ExCB or the ExTL involved remains fully responsible for the test results and the content of the test report.

### A “Witness Testing Agreement” shall be signed between the ExCB or the ExTL, and the manufacturer or user test facility. The agreement shall cover the following, as a minimum:

- The scope of work (specific tests or test standards) for which the participating test facility is qualified to carry out.

- The facilities and services provided by the manufacturer or user test facility

- The right of initial assessment and re-assessment of the manufacturer or user test facility by the ExTL or the ExCB.

- The supervisory role of ExTL or ExCB staff witnessing the tests.

- Use of data generated.

### The IECEx Secretary shall maintain and publish a Register of participating manufacturers and end users, together with the associated Test Standard(s) for which they are qualified.

## Role and Responsibilities of the ExTL

## Where an ExTL is involved, it shall be responsible for the following:

### Have the required quality procedures which adequately cover all aspects of witness testing.

### Ensure that its staff involved in witness testing are:

- Qualified for the test standard involved.

- Provided with the necessary initial and on-going training on the requirements of this Operational Document.

### Carry out the required initial and on-going site assessments of the participating test facility

### Ensure the presence of ExTL staff at the participating test facility during testing to witness all aspects of the tests carried out by personnel of the participating test facility, except as permitted under clause 6.9 and 6.10.

### Check the competence of personnel assigned to conduct testing on behalf of the participating test facility to perform the required tests.

### Review and sign the required test report.

### Sign the agreement referenced in 6.1.2 above.

## Role and Responsibilities of the ExCB

The ExCB shall be responsible for the following activities:

### Ensuring that its staff involved in witness testing are

- Qualified for the test standard involved.

- Provided with the necessary initial and on-going training on the requirements of this procedure.

### Where an ExTL is involved in witness testing:

- Assess the quality procedures of the ExTL applicable to the operation of witness testing.

- Review of documentation held by the ExTL covering site assessments/reassessments of the participating test facility, and contractual arrangements.

- Ensure there is an appropriate three-way agreement between the ExCB, the ExTL, and the participating test facility, as per 6.1.2 above.

### Where an ExTL is not involved in witness testing, the ExCB must:

- Have its own appropriate quality procedures and documentation for the operation of a witness testing program.

- Carry out the required initial and on-going site assessments and reassessments of the participating test facility.

- Check the competence of personnel assigned to conduct testing on behalf of the participating test facility to perform the required tests.

- Ensure the presence of qualified ExCB staff at the participating test facility to witness all aspects of the tests carried out by personnel of the participating test facility.

- Review and sign the test report.

- Ensure there is an appropriate two-way agreement as per 6.1.2 above.

### Registration of the manufacturer or user test facility with the IECEx Secretariat and updating of the current information in the Register.

### Issuing ExTR and IECEx Certificates based on witness testing.

## Role and Responsibilities of the Manufacturer or User Test Facility

Participating manufacturer or user test facilities are responsible for the following:

### Make a formal application to any ExTL or ExCB to participate in the witness testing program.

### Identify the scope of testing it wishes to conduct at its facility.

### Sign the required agreement with the ExTL, and/or the ExCB involved.

### Demonstrate that the facility is in compliance with all the relevant requirements of ISO/IEC 17025 and the requirements of the witness testing program. Relevant requirements of ISO/IEC 17025 shall include the following elements, as a minimum:

- Critical consumables

- Personnel

- Accommodations and environmental conditions

- Equipment

- Measurement traceability / calibration

- Identification of test items

### Appoint appropriate personnel to be responsible for the witness testing program at the participating test facility.

### Ensure that all tests are carried out in accordance with the applicable test standard(s), instructions given by the ExCB or ExTL staff witnessing the test, and all the relevant provisions of the agreement.

### Sign the test report as required.

### Participate in proficiency testing as directed by the ExCB involved.

## Choice of ExCBs/ExTLs by the Manufacturer or User Test Facility

### A participating test facility shall be permitted to use more than one ExCB or ExTL for a given test standard (i.e. have an agreement with more than on ExCB).

### A participating test facility shall not be permitted to use more than one ExCB or ExTL for the same product model.

Note - a participating test facility shall be permitted to address a test report for the same model to more than one ExCB/ExTL for the purpose of issuing a CoC.

## Site Assessment/Reassessment of the Manufacturer or User Test Facility

### Qualified ExTL or ExCB staff shall carry out site assessments of the participating test facility in accordance with own internal quality procedures.

### The initial assessment shall be performed after a formal application is received from the participating test facility and prior to conducting actual witness testing.

### The continued validity of the initial assessment or reassessment shall be evaluated during each subsequent witness testing project/visit. The evaluation shall consider the following factors:

- The scope of work defined in the agreement referenced in clause 6.1.2

- The frequency of visits

- The extent of organizational or equipment changes at the facility

- Other changes that affect the performance of the participating test facility relative to the conduct of off-site testing.

### The site assessment shall cover relevant requirements of ISO/IEC 17025. Relevant requirements of ISO/IEC 17025 shall include the following elements, as a minimum:

- Critical Consumables

- Control of environmental conditions at the participating facility.

- Stability and capacity of power supplies.

- Availability and adequacy of test equipment.

- Measurement traceability/calibration.

- Competence of personnel at the participating test facility assigned to conduct the witness testing activities.

- Identification of test items.

### The ExCB shall document the results of any assessment/reassessment activity, using the appropriate IEXEx Site Assessment Checklist (Annex A). Where the ExTL is involved, it shall provide a copy of the assessment/reassessment report to the ExCB for review and approval.

## Requirements for Conducting Tests and Issuing Test Reports and Certificates

### Prior to commencing the tests, witnessing staff of the ExTL or ExCB involved shall:

- Prepare the necessary test plan and review it with personnel of the participating test facility assigned to perform the tests.

- Check the test set-up for correctness.

- Check that the appropriate test instruments are used and ensure their accuracy and proper functioning.

- Verify that the test instruments are properly calibrated by an ISO 17025 accredited calibration lab and that calibration is traceable to National Standards.

- Provide the necessary work instructions and direction.

- Verify that the test sample(s) is/are representative of production and properly and uniquely identified.  In cases where multiple samples are required for testing, they are to be identifiable to the testing for which they were subjected.

### Witnessing staff of the ExCB or ExTL involved shall be present during testing and shall continue to supervise and check all critical aspects of the tests.

### Test reports/records prepared shall meet the following criteria:

- Be the relevant IECEx Test Report Forms

- Include a unique identification number to permit correlation to the ExTR.

- Record the date of testing and identification of all the test samples.

- Indicate the name and address of the manufacturer or user test facility used.

- Indicate what tests have been carried out at the manufacturer or user test facility.

- Record all test instruments used in testing, indicating calibration status.

- Include applicable test data.

- Show the following statement: “The test facility was deemed to have the environment and the capabilities necessary to perform the tests indicated in the test report”.

- Be signed by: the person who carried out the testing on behalf of the participating test facility under “Tested by”; the person who witnessed the tests on behalf of the ExCB or ExTL involved under “Witnessed by”; and the person who reviewed the report, other than the witnessing staff under “Authorized by”.

### When an ExTR is based on witness testing, the following shall apply:

### A reference to the use of the manufacturer or user test facility shall be included as “Additional Information” in the on-line ExTR summary.

### All documentation related to assessment/reassessment activities shall be included as a permanent attachment to the ExTR

## Long term tests

### Long term tests are not required to be witnessed for the entire duration of the test. Such tests shall be witnessed at the start and at the conclusion of the test.

### The test sample(s) shall be identified prior to the start of the test and witnessed by ExTL or ExCB staff.

### The test set-up shall be subject to continuous monitoring of conditions.

### The conclusion of the test shall be witnessed by ExTL or ExCB staff.

### Records of the continuous monitoring system shall be reviewed by the ExTL.

### The start and conclusion of tests do not need to be witnessed by the same member of the ExTL or ExCB staff.

## Remote Witnessing of Testing

### Remote witness testing may be used only where, based on past experience and prior site assessment, ExTL or ExCB staff have confidence in the ability of the persons performing the tests and the test equipment used. Confidence is based on he the following conditions:

- The manufacturer shall have had a minimum of three witnessed tests by an ExTL or ExCB, with satisfactory results.

- The ExTL or ExCB shall have a process to demonstrate that the necessary level of trust and confidence is maintained through periodic witnessing of testing by the ExTL or ExCB.

### Where remote witness testing is carried out, the following guidelines shall be followed:

- All test equipment calibration records requested by the witnessing entity shall be made available and correlated with the designations shown on the test equipment used.

- Details of the test plan and test set-up (including for instance placement of thermocouples, required photos, etc.) and testing date shall be discussed and agreed upon prior to commencing testing.

- Prior to commencing testing, detailed photos of the test set-up shall be sent to the ExTL or ExCB for review and approval. If necessary, the witnessing entity may require that “live” viewing of the test set-up using a video camera be available to check the set-up prior to testing.

- Acceptance of tests and related data witnessed shall be determined by ExTL or ExCB staff.

- The required agreement between the participating test facility and the ExTL and/or ExCB shall include a stipulation confirming that:

- The condition of clause (d) above is understood.

- Witnessing of tests does not imply acceptance of the data, if, in the opinion of the ExTL or the ExCB staff, the complexity of the tests or the results obtained may require further evaluation.

- Test samples used shall be distinctly marked.

- The ExTL or ExCB may request that test samples are provided for additional review.

## Partially Witnessed Testing

### Partially witnessed testing is a common practice currently used by ExTLs and ExCBs whereby witnessing only some parts of the agreed upon testing program is deemed sufficient. Examples of situations where partially witnessed testing is used include the following:

- Long term tests as described in 6.8 above

- Remote witnessing of testing as described in 6.9 above

- Other scenarios where the ExTL or ExCB has gained and maintains a high level of trust and confidence in those conducting the tests, based on past experience.

### Where partially witnessed testing is used under scenario (c) above, the following shall be adhered to:

- The manufacturer shall have had a minimum of three witnessed tests by an ExTL or ExCB, with satisfactory results.

- The ExTL or ExCB shall have a process to demonstrate that the necessary level of trust and confidence is maintained through periodic witnessing of testing by the ExTL or ExCB.

- Details of the test plan and test set-up shall be reviewed and approved by ExTL or ExCB staff prior to commencing testing.

- The transition from witness testing to partially witnessed testing shall be subject to acceptance by the ExCB in question.

### Where the manufacturer uses more than one ExTL or ExCB for partially witnessed testing, each ExTL or ExCB may accept the confidence building measures gained by other ExTL or ExCB.