IECEx Guide on Uncertainty of Measurement

Tests and measurements required in the Standards IEC 60079-series

(This draft results from practice in Germany and the experience of Exam/BVS in Dortmund and PTB in Braunschweig)

In the tables (1) means:

Measuring uncertainty of general industrial standard is sufficient, failure of the measuring devices will be detected easily. The measuring uncertainty of the devices has to provide the safe estimation of the values required in the standards taking into account the stated tolerances. Where no tolerances are given in the standard, a measuring uncertainty of 5% shall be adequate.

The tables cover up to now only IEC 60079-0, 60079-5, 60079-6 , IEC 60079-7, 60079-11 and 60079-18.

| SPECIFIC TEST ITEMS |  |  |  |  |  | Sheet 1 (2) |
| --- | --- | --- | --- | --- | --- | --- |
| Electrical apparatus for explosive gas atmospheres General requirements IEC 60079-0:1998 |  |  |  |  |  |  |
| 1 | 2 | 3 | 4 | 5 | 6 | 7 |
| Clause | Requirement | Kind of test, measurand where appropriate | Test equipment | Standard | Cali-bration | Remarks |
| 23.4.3.1 23.4.3.2 | resistance to impact test drop test | - length - mass | folding rule (cf. Annex D) weighing instru-ment (cf. Annex D) |  | no | no special measurement requirements from the safety view-point (1) |
| 23.4.4 | degree of protection IP of enclosures | ingress of foreign matter and water | test bodies, dust test chamber, spray head or flushing pipe | IEC 60529 | no |  |
| 23.4.5 | torque test for bushings | torque measurement | torque wrench |  | no | (1) |
| 23.4.6 | temperature rise test | - electric power - tempera- ture - time | current, voltage, power usual electro-technical equipment various measuring principles are possible stop watch |  | yes  1% yes  5 K no | Measurement uncertainty U shall not be added to result when type testing (U covered by 5K/10K increase required by standard) |

| SPECIFIC TEST ITEMS |  |  |  |  |  | Sheet 2 (2) |
| --- | --- | --- | --- | --- | --- | --- |
| Electrical apparatus for explosive gas atmospheres General requirements IEC 60079-0:1998 |  |  |  |  |  |  |
| 1 | 2 | 3 | 4 | 5 | 6 | 7 |
| Clause | Requirement | Kind of test, measurand where appropriate | Test equipment | Standard | Cali-bration | Remarks |
| 23.4.7 Nos. 1-4 | non-metallic enclosures, aging stability | - cold - heat - humidity | climatic cabinet |  | yes | tolerances specified in IEC 60079-0 |
| 23.4.7.5 | resistance to light |  | cf. standards ISO 4892-1 ISO 179 |  | no |  |
| 23.4.7.6 | resistance to chemical agents |  | liquids acc. to 23.4.7.6 ISO 1817 |  | no |  |
| 23.4.7.7 cf. 23.4.3 |  |  |  |  |  |  |
| 23.4.7.8 | static electricity | insulation resistance | cf. IEC 60079-0 23.4.7.8 |  | no | (1) |

| SPECIFIC TEST ITEMS |  |  |  |  |  | Sheet 1 (1) |
| --- | --- | --- | --- | --- | --- | --- |
| Electrical apparatus for explosive gas atmospheres Oil-immersion "o" IEC 60079-6:1995 (Ed.2) |  |  |  |  |  |  |
| 1 | 2 | 3 | 4 | 5 | 6 | 7 |
| Clause | Requirement | Kind of test, measurand where appropriate | Test equipment | Standard | Cali-bration | Remarks |
| 5.1 | Sealing of enclosures | pressure test IP test | pressure gauge cf. IEC 60079-0 | IEC 60079-6 | yes  5% |  |
| 5.2 | leakage rate < 5% | reduced pressure test | pressure gauge clock | IEC 60079-6 | yes  1% no |  |

| SPECIFIC TEST ITEMS |  |  |  |  |  |  | Sheet 1 (2) |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Electrical apparatus for explosive gas atmospheres Powder filling "q" IEC 60079-5:1997 (Ed. 2) |  |  |  |  |  |  |  |  |
| 1 | 2 | 3 |  | 4 | 5 | 6 |  | 7 |
| Clause | Requirement | Kind of test, measurand where appropriate |  | Test equipment | Standard | Cali-bration |  | Remarks |
| 5.1.1 | mechanical strength pressure type test |  | resistance to impact test (IEC 60079-0, 23.4.3.1) pressure type test of enclosure (5.1.1) - pressure - length - time | cf. IEC 60079-0 pressure measuring device length measuring device clock |  | yes ± 5% no no |  | (1) (1) |
| 5.1.2 | degree of protection of enclosure |  | test of degree of protection of enclosure - IP degree of protec- tion (IEC 60070-0, No. 23.4.4) | cf. IEC 60079-0 | IEC 60529 |  |  |  |
| SPECIFIC TEST ITEMS |  |  |  |  |  |  | Sheet 2 (2) |  |
| Electrical apparatus for explosive gas atmospheres Powder filling "q" IEC 60079-5:1997 (Ed.2) |  |  |  |  |  |  |  |  |
| 1 | 2 | 3 |  | 4 | 5 | 6 |  | 7 |
| Clause | Requirement | Kind of test, measurand where appropriate |  | Test equipment | Standard | Cali-bration |  | Remarks |
| 5.1.4 | electric strength test |  | electric strength test of filling material - voltage - current |  |  | yes  1% yes  1% |  |  |
| 5.1.3 | flammability of materials |  | flamma-bility |  | cf. IEC 60079-1 | ISO 1210 or UL 94 VO or IEC 707, FV method |  |  |

| SPECIFIC TEST ITEMS |  |  |  |  |  |  | Sheet 1 (5) |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Electrical apparatus for explosive gas atmospheres Intrinsic Safety "i" IEC 60079-11:1999 (Ed. 4) |  |  |  |  |  |  |  |  |
| 1 | 2 | 3 |  | 4 | 5 | 6 |  | 7 |
| Clause | Requirement | Kind of test, measurand where appropriate |  | Test equipment | Standard | Cali-bration |  | Remarks |
| 6.2 8.7 6.3 6.4 | Requirements as regards dimensions for the construction of components and apparatus, e.g.: |  | length determina-tion of: lengths, cross-sections, thicknesses and widths of conduct-ing tracks, surfaces, creepage distances and clearances, separation distances, partition thicknesses | length measuring device |  | no |  | Examined are: drawings, foils, samples (1) |
| 10.1 to 10.3 | spark ignition test |  | ignition test | spark-test apparatus gas mixing device | IEC 60079-11, 10.1 and Annex B | yes, cf. 10.3 yes |  | test gas must be of sufficient quality (1) |

| SPECIFIC TEST ITEMS |  |  |  |  |  |  | Sheet 2 (5) |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Electrical apparatus for explosive gas atmospheres Intrinsic Safety "i" IEC 60079-11:1999 (Ed 4) |  |  |  |  |  |  |  |  |
| 1 | 2 | 3 |  | 4 | 5 | 6 |  | 7 |
| Clause | Requirement | Kind of test, measurand where appropriate |  | Test equipment | Standard | Cali-bration |  | Remarks |
| 10.5 6.2.4 8.1.4 8.2 10.9.3 | temperature tests, temperature rise measurement determination of the max. surface temperature of cells and batteries |  | temperature measurement temperature rise test upon short-circuit | all conventional temperature measuring instruments all conventional temperature measuring instruments | IEC 60079-0 | see IEC 60079-0 |  | Influence of the measuring element on the temperature prevailing at the measuring point shall be largely excluded. |
| 10.6, 6.4.12, 5.4, 6.4.4, 6.4.5, 6.4.11, 8.1.3, 8.1.4, 8.2, 8.5, 8.8 | insulation voltage test, electric strength |  | voltage current | a.c. or d.c. voltage source, voltage detector, ammeter | HD 401S1 IEC 60079-11, 10.6 | yes  2% yes  2% |  | AC 48 up to 62 Hz, DC ripple voltage max. 3% |
| 10.7 6.2.4 | small component ignition test |  | ignition test – no ignition due to small hot component | explosion chamber test mixture gas mixing device | IEC 60079-11, 10.7 | no yes, special test |  | test gas must be of sufficient quality (1) |

| SPECIFIC TEST ITEMS |  |  |  |  |  |  | Sheet 3 (5) |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Electrical apparatus for explosive gas atmospheres Intrinsic Safety "i" IEC 60079-11:1999 (Ed.4) |  |  |  |  |  |  |  |  |
| 1 | 2 | 3 |  | 4 | 5 | 6 |  | 7 |
| Clause | Requirement | Kind of test, measurand where appropriate |  | Test equipment | Standard | Cali-bration |  | Remarks |
| 10.9 7.4 | test for cells and batteries |  | determina-tion of open circuit voltage, determina-tion of internal resistance - voltage - current | voltmeter voltmeter ammeter |  | yes  2% yes  2% yes  2% |  |  |
| 10.9 | including spark ignition test |  | ignition test | spark-test apparatus | IEC 60079-11, 10.2 and Annex B | yes, 10.3 |  |  |
| 10.10.1 6.4.4, 6.7, 8.4, 8.6, 9.2.2 | mechanical strength a)casting compounds: strength strength of free surfaces |  | force impact test | metal rod impact test apparatus | IEC 60079-11, 10.10.1 IEC 60079-0 | no no |  | (1) |

| SPECIFIC TEST ITEMS |  |  |  |  |  |  | Sheet 4 (5) |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Electrical apparatus for explosive gas atmospheres Intrinsic Safety "i" IEC 60079-11:1999 (Ed.4) |  |  |  |  |  |  |  |  |
| 1 | 2 | 3 |  | 4 | 5 | 6 |  | 7 |
| Clause | Requirement | Kind of test, measurand where appropriate |  | Test equipment | Standard | Cali-bration |  | Remarks |
| 10.10.2 6.3, 6.4 | b) partitions: strength |  | force | massive test rod | IEC 60079-11, 10.10.2 | no |  | (1) |
| 7.4.7 | c) drop test for portable electrical apparatus with batteries |  | drop test |  | IEC 60079-0 23.4.3.2 | no |  |  |
| 10.11 7.7 | d) for apparatus containing piezoelectric devices determination of the capacitance and voltage during impact test |  | impact test voltage capacitance | impact test apparatus voltmeter capacitance meter | IEC 60079-0 | no yes  2% yes  2% |  |  |
| 10.13 | e) cable pull test |  | pull test, force |  | IEC 60079-11, 10.13 | no |  | (1) |

| SPECIFIC TEST ITEMS |  |  |  |  |  |  | Sheet 5 (5) |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Electrical apparatus for explosive gas atmospheres Intrinsic Safety "i" IEC 60079-11:1999 (Ed.4) |  |  |  |  |  |  |  |  |
| 1 | 2 | 3 |  | 4 | 5 | 6 |  | 7 |
| Clause | Requirement | Kind of test, measurand where appropriate |  | Test equipment | Standard | Cali-bration |  | Remarks |
| 10.12 7.3, 8.6 and 9 | determination of the characteris-tics of current- voltage transients type tests for diode safety barriers and safety shunts |  | voltage current time | pulsed current source, storage oscilloscope, current source, voltmeter |  | no no no yes  2% |  | (1) (1) (1) |

| SPECIFIC TEST ITEMS |  |  |  |  |  |  | Sheet 1 (5) |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Electrical apparatus for explosive gas atmospheres Increased Safety "e" IEC 60079-7:2001 (Ed. 3) |  |  |  |  |  |  |  |  |
| 1 | 2 | 3 |  | 4 | 5 | 6 |  | 7 |
| Clause | Requirement | Kind of test, measurand where appropriate |  | Test equipment | Standard | Cali-bration |  | Remarks |
| 4.2 | terminals for external connections |  | force temperature | torque screwdriver force measurement or pull weights millivoltmeter or recorder with thermocouples |  | no no yes  5 K |  |  |
| 4.3 | internal connections |  | force temperature | torque screwdriver force measurement or pull weights millivoltmeter or recorder with thermocouples |  | no no yes  5 K |  |  |
| 4.4 | clearances |  | length measurement | probes gauges rules, caliper gauges, measuring magnifier | IEC 60664-1 | no |  |  |
| 4.5 | creepage distances |  | length measurement | probes gauges rules, caliper gauges, measuring magnifier | IEC 60664-1 | no |  |  |

| SPECIFIC TEST ITEMS |  |  |  |  |  |  | Sheet 2 (5) |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Electrical apparatus for explosive gas atmospheres Increased Safety "e" IEC 60079-7:2001 (Ed.3) |  |  |  |  |  |  |  |  |
| 1 | 2 | 3 |  | 4 | 5 | 6 |  | 7 |
| Clause | Requirement | Kind of test, measurand where appropriate |  | Test equipment | Standard | Cali-bration |  | Remarks |
| 4.6 | solid electrical insulating materials |  | mechanical stability | climatic chamber or warming cabinet |  |  |  | cf. IEC 60079-0 |
| 4.7 | windings |  | voltage | voltage source voltmeter | IEC 60251-1 | yes 2% |  |  |
| 4.8 | limiting temperature |  | temperature excess temperature | facilities for applying thermal load ammeter voltmeter resistance measuring ranges thermocouples + millivoltmeters |  | cf. IEC 60079-0 |  |  |
| 4.10 | degrees of protection provided by enclosures |  | ingress of foreign bodies and water | test bodies, dust test chamber spray head or flushing pipe | EN 60529 |  |  | cf. IEC 60079-0 |

| SPECIFIC TEST ITEMS |  |  |  |  |  |  | Sheet 3 (5) |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Electrical apparatus for explosive gas atmospheres Increased Safety "e" IEC 60079-7:2001 (Ed.3) |  |  |  |  |  |  |  |  |
| 1 | 2 | 3 |  | 4 | 5 | 6 |  | 7 |
| Clause | Requirement | Kind of test, measurand where appropriate |  | Test equipment | Standard | Cali-bration |  | Remarks |
| 5.2 6.3 | rotating electrical machines |  | air gap temperature rise with rotors locked | gauges, feeler gauges, measuring tool for internal and external measurement voltage supply, voltage measurement, current measurement (a.c.), resistance measurement, thermocouples and millivolt recorders |  | no yes  0,5% yes  0,5% yes  0,2% yes  2 K |  |  |
| 5.3, 6.3 | luminaires designed for mains supply |  | distance temperature contact force | probe or gauge voltage supply voltmeter thermocouple and millivoltmeter force measurement or pull weights |  | no yes  0,5% yes  2 K no |  |  |

| SPECIFIC TEST ITEMS |  |  |  |  |  |  | Sheet 4 (5) |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Electrical apparatus for explosive gas atmospheres Increased Safety "e" IEC 60079-7:2001 (Ed.3) |  |  |  |  |  |  |  |  |
| 1 | 2 | 3 |  | 4 | 5 | 6 |  | 7 |
| Clause | Requirement | Kind of test, measurand where appropriate |  | Test equipment | Standard | Cali-bration |  | Remarks |
| 6.6.2 | batteries |  | insulation resistance | ohmmeter |  | yes  5% |  |  |
| 6.6.3 | mechanical strength of battery cells |  | shock test capacitance measurement - current - voltage - time | shock machine charging/discharg-ing device | IEC 60068-2-27 | yes  1% yes  1% yes  1% |  |  |
| 6.6.4 | adequate venti-lation of battery containers |  | gas measurement - hydrogen -temperature | charging/discharg-ing device test room free from draught |  | yes  0,1% percent by volume H2 yes  5 K |  |  |

| SPECIFIC TEST ITEMS |  |  |  |  |  |  | Sheet 5 (5) |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Electrical apparatus for explosive gas atmospheres Increased Safety "e" IEC 60079-7:2001 |  |  |  |  |  |  |  |  |
| 1 | 2 | 3 |  | 4 | 5 | 6 |  | 7 |
| Clause | Requirement | Kind of test, measurand where appropriate |  | Test equipment | Standard | Cali-bration |  | Remarks |
| 6.8 | resistance heating units |  | resistance to impact test crushing test low tempera-ture bend test limiting temperature | rig for resistance to impact test test apparatus, force measuring device or weighing instrument, resistance measurement test apparatus, resistance measurement voltmeter, thermocouple and millivoltmeter | IEC 60079-0 | yes  2% yes  5% yes  5% |  | cf. IEC 60079-0 cf. IEC 60079-0 cf. IEC 60079-0 cf. IEC 60079-0 |

| SPECIFIC TEST ITEMS |  |  |  |  |  |  | Sheet 1 (x) |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Electrical apparatus for explosive gas atmospheres Encapsulation "m" IEC 60079-18:1992 (ed. 1) |  |  |  |  |  |  |  |  |
| 1 | 2 | 3 |  | 4 | 5 | 6 |  | 7 |
| Clause | Requirement | Kind of test, measurand where appropriate |  | Test equipment | Standard | Cali-bration |  | Remarks |
| 5.4 6.1 | prevention of excessive tempe-rature rise by determination of the max. temperature of primary and secondary cells, batteries and accumulators |  | temperature | temperature measuring device |  |  |  | cf. IEC 60079-0 |
| 6.2 | maximum tempera-ture upon ruptur-ing of the fuse must not damage the compound to such an extent that the type of protection is affected |  | temperature | temperature measuring device |  |  |  | cf. IEC 60079-0 |
| 8.1.1 | suitability of compound |  | electric strength | high-voltage tester, warming cabinet |  | yes |  | no flashover or breakdown may occur |

| SPECIFIC TEST ITEMS |  |  |  |  |  |  | Sheet 2 (x) |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Electrical apparatus for explosive gas atmospheres Encapsulation "m" IEC 60079-18:1992 |  |  |  |  |  |  |  |  |
| 1 | 2 | 3 |  | 4 | 5 | 6 |  | 7 |
| Clause | Requirement | Kind of test, measurand where appropriate |  | Test equipment | Standard | Cali-bration |  | Remarks |
| 8.1.2 | suitability of compound |  | water absorption | weighing instrument (balance) |  | yes  0,1% |  | increase in mass not more than 1% |
| 8.2.1 8.2.1.1 | resistance to thermal load |  | maximum temperature | temperature probe, temperature measuring device |  |  |  | cf. IEC 60079-0 |
| 8.2.1.2 |  |  | thermal cycling test | warming cabinet with temperature control and possibility of external power supply of unit or the like; voltmeter and/or ammeter, resistance measuring device, wattmeter |  | yes  5K yes  1% yes  1% yes  2% |  | continuous recording of test conditions and of measurement results; acceptance criteria acc. to 6.2.1.3 |
| 8.2.2 | Cable pull test |  | mechanical test | Application of tensile force by weight |  | Yes  5% |  |  |

| SPECIFIC TEST ITEMS |  |  |  |  |  |  | Sheet 3 (x) |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Electrical apparatus for explosive gas atmospheres Encapsulation "m" IEC 60079-18:1992 |  |  |  |  |  |  |  |  |
| 1 | 2 | 3 |  | 4 | 5 | 6 |  | 7 |
| Clause | Requirement | Kind of test, measurand where appropriate |  | Test equipment | Standard | Cali-bration |  | Remarks |
| 8.2.3 and 4.5.3 | isolation |  | Electric strength test | high-voltage tester, warming cabinet |  | yes |  |  |