### INTERNATIONAL ELECTROTECHNICAL COMMISSION (IEC) SYSTEM

FOR CERTIFICATION TO STANDARDS RELATING TO EQUIPMENT FORUSE

IN EXPLOSIVE ATMOSPHERES (IECEx SYSTEM)

Title: Report from Dr Uwe Klausmeyer, Convenor ExTAG WG 10, Proficiency Testing

Circulated to: Ex Testing and Assessment Group (ExTAG)

INTRODUCTION

This document is an updated report prepared by Dr Uwe Klausmeyer the Convenor of ExTAG WG10, Proficiency Testing and replaces previous ExTAG/236/R.  This Report will be presented by Dr. Klausmeyer during the ExTAG Split Meeting for discussion and consideration.

Michel Brenon

ExTAG Secretary

| Address: Standards Australia Building 286 Sussex Street Sydney NSW 2000 Australia | IECEx Secretariat Contact Details: Tel: +61 2 8206 6940 Fax: +61 2 8206 6272 e-mail: chris.agius@iecex.com http://www.iecex.com | ExTAG Secretary Mr Michel Brenon/LCIE Tel: +33 1 40 95 5519 Fax: +33 1 40 95 5520 e-mail michel.brenon@lcie.fr |
| --- | --- | --- |

Report of the convenor of ExTAG WG 10 “Proficiency Testing”, Dr. Klausmeyer, for the ExTAG meeting in Split 2011

Updated Progress report

Two years of PTB Ex Proficiency Testing Program (Pilot)

PTB Ex Proficiency Testing Team continued to establish the system for interlaboratory comparisons in explosion protection with two engineers (Jia Wu – Test Round “i”, Tim Krause – Test Round “d”) supported by experts of the involved laboratories. Because of the unexpected high number of participating laboratories (38 for “d”, 37 for “i”, see the list at the end of this report), the pilot programs needed a longer time for launching. From the beginning of the program the team assisted the participants in all kinds of questions and problems inclusive on site visits, if requested. Up to date all problems with the test samples could be resolved by correspondence. First results have been received by the laboratories.

Status of the ongoing development of the Ex PTP

The test samples for all participating laboratories (both test rounds) have been developed, produced, prepared and made available until June (“d”) and August (“i”) 2011. The identical copies have been submitted to homogeneity and stability tests. The “i” Test Boxes have been submitted to burn in tests under elevated temperature in order to check the reliability of the electronic components. Each box has been sealed by a sticker in order to avoid not authorized manipulation. During 2011 a program for data analysis and evaluation has been developed according to the requirements of ISO/IEC 17043. The statistical models and tools are based on ISO 13528.They have been developed for professional evaluation and proper presentation of the test results.

Tendency of first results received until 01st of July, 2011

Test Round “d”: Almost 50% of the participating laboratories have uploaded their results. The evaluated average of the explosion pressure for a certain test set up and gas mixture in the small vessel has been calculated statistically for each laboratory. According to ISO 13528 the assigned value has been determined as a robust average of the results reported by the participants. All results are confidential until the pilot program is finished. The received data show some differences and are a substantial basis for the discussion how to reduce the range; after laboratories get knowledge of their individual results (phase II) PTB will assist to review the individual measurement methods and to calculate an accuracy of their submitted results which will be included in the final evaluation report.

Test Round “i”: 25% of the laboratories have uploaded their results. Up to now the average values of the “number of contacts until ignition” of four different circuits have been calculated. The average values of all participating laboratories determined the so-called assigned value. The relative difference between the average value of each measurement series of the participating laboratories and the assigned value was significant related to the assigned value. The test results should be discussed in phase II with the intention to minimize the range and to analyse all results together with the participating laboratories in order to improve the reproducibility.

Conclusions and possible action in reflecting first results

It is now evident that an Ex PTP is feasible to conduct and brings valuable knowledge for the participating laboratories. The confidence of customers and accreditation bodies is directly influenced positively in being aware of the Ex PTP participation. Especially the IECEx system is gaining credibility by the fact that 86% of the participating laboratories are IECEx ExTLs. Both, proficiency testing and quality assessment of a laboratory provide evidence of competence. Both elements could be connected e.g. by performing an explosion pressure measurement in the Test Sample “d” during the onsite assessment. The technical guidance documents could be edited in this sense. Consequence of this approach could be that those laboratories not able to perform these tests properly could run in trouble concerning their scope. The feed back loop proficiency testing and quality assessment resp. scope of a laboratory could be a powerful source for quality improvement (CIP). In this context it is necessary to decide on future PTPs (e.g. electrostatic charge measurements, temperature rise tests) in using a comprehensive matrix of all tests done under the scope of IECEx. ExTAG WG 10 will take all these aspects on board and draft decisions for ExTAG to be discussed and passed on to ExMC for endorsement.

Main agenda items for the future

- 05th of September, 2011 in Split: meeting of the ExTAG WG 10 (venue: room Vis, Le Meridian Lav, agenda: review of the Ex PTP and the first results, conclusions on the results to be discussed like “Ex PTP mandatory for each ExTL or not”, interaction Ex PTP and on site assessment of the ExTLs, future PTP programs like electrostatic charge tests or temperature rise tests, recommendations for the ExTAG meeting 06th of September 2011)

- January 2012: phase II of the programs

- May 2012: final results to be compiled for the participating laboratories

- June 2012: workshop at PTB organized by the Ex PTP team for all participating laboratories (agenda example: how an explosion pressure can be measured, how to improve uncertainty of measurement)

- September 2012 (annual meeting of IECEx): 
- ExTAG WG 10 meeting: proposing that ExMC may endorse the Ex PTP as mandatory for all ExTLs (via ExTAG), development of an organization and finance model which includes a review of the IECEE practice

| Test laboratory | Country |
| --- | --- |
| Baseefa | United Kingdom |
| Bundesanstalt für Materialforschung und -prüfung | Germany |
| Bureau Veritas Consumer Product Services GmbH | Germany |
| Central Mining Institute Experimental Mine "BARBARA" | Poland |
| CESI Centro Electtrotecnico Sperimentale Italiano S.p.A. | Italy |
| China National Quality Supervision & Test Center for Explosion-Proof and Safety Products für Coal Mine (CMExC) | China |
| China National Quality Supervision and Test Center for Explosion Protected Electrical Apparatus (CQST) | China |
| CSA International | Canada |
| DEKRA EXAM GmbH | Germany |
| Det Norske Veritas Certification AS | Norway |
| DVLA - LABEX (Laboratory of explosion proof test) | Brasil |
| ELECTRONICS REGIONAL TEST LABORATORY (EAST) | India |
| Ex-Agencija Laboratory for explosion protection ExLAB | Croatia |
| FM Approvals, LLC | USA |
| IMQ S.p.A. | Italy |
| INERIS | France |
| Institut für Sicherheitstechnik IBE | Germany |
| International Test and Cerification (ITACS) | Australia |
| Intertek | USA |
| Intertek Testing & Certification Ltd | United Kingdom |
| ISSeP | Belgium |
| KEMA | Netherlands |
| Korea Gas Safety Corporation (KGS) | Korea |
| LABELO / PUCRS | Brasil |
| Mine Safety Technology Centre | Australia |
| National Supervision and Inspection Center for Explosion Protection and Safety Instrumentation (NEPSI) | China |
| Physikalisch Technische Bundesanstalt (PTB) | Germany |
| Physical Technical Testing Institute (FTZÚ) | Czech Republic |
| Supervision & Test Centre of Ex-products of China Petroleum & Chemical Industry (PCEC) | China |
| Simtars | Australia |
| SIQ Ljubljana | Slovenia |
| Sira Test & Certification Ltd | United Kingdom |
| Testing Laboratory of ExVÁ Ltd. | Hungary |
| TestSafe | Australia |
| TL CCVE (NANIO "CCVE") | Russia |
| TÜV AUSTRIA SERVICES GMBH | Austria |
| TÜV NORD CERT GmbH | Germany |
| UkrNIIVE, Testing Certification Center of Explosion protected and mining electrical Equipment | Ukraine |
| UL Demko | Denmark |
| Underwriters Laboratories Inc. | USA |
| VTT Expert Services Ltd | Finland |
| Zelm Explosionsschutz GmbH | Germany |