INTERNATIONAL ELECTROTECHNICAL COMMISSION IEC SYSTEM FOR

CERTIFICATION TO STANDARDS RELATING TO EQUIPMENT FOR USE IN

EXPLOSIVE ATMOSPHERES (IECEx SYSTEM)

TITLE: Re-assessment and Scope Extension Report  for the continued 	acceptance of Korea Gas Safety Corporation, KGS,  an Accepted 	Certification Body 	(ExCB), and Accepted Ex Testing Laboratory (ExTL) 	within the IECEx System, Equipment Scheme 02.

Circulation to: Members of the IECEx Management Committee, ExMC

INTRODUCTION

In accordance with the 5 year re-assessment plan for the surveillance and monitoring of bodies within the IECEx System, this document contains the IECEx Re-assessment and Scope Extension Report for the continued acceptance of Korea Gas Safety Corporation, KGS, as an Accepted Certification Body(ExCB) and an Accepted Ex Testing Laboratory(ExTL) within the IECEx System, Equipment Scheme 02.

During the re-assessment the team also conducted a review and assessment of the ExCB’s capability to include the following standards within its scope.-

ISO 80079-36 Edition 1	Explosive atmospheres – Part 36: Non-electrical 	Equipment for explosive atmospheres - Basic method and requirements

ISO 80079-37 Edition 1	Explosive atmospheres – Part 37: Non-electrical 	Equipment for explosive atmospheres – Non-electrical type of protection 	constructional safety “c”, control of ignition source “b”, liquid immersion “k”

This report covers both the re-assessment and the scope extension for IEC 80079-36 and 37.

Chris Agius

IECEx Secretariat

| IECEx Secretariat Australia Square Level 33, 264 George Street Sydney NSW 2000 Australia | Tel: +61 2 4628 4690 Fax: +61 2 46 27 5285 Email: info@iecex.com |
| --- | --- |

IEC System for certification to standards relating to equipment for use in Explosive Atmospheres (IECEx System)

IECEx Assessment Report Form

IECEx Assessment Report Form for use by IECEx Assessment Teams to report Assessments conducted according to the IECEx Assessment Procedures of

- Operational Document IECEx OD003-2 for the Certified Equipment Scheme

- Operational Document IECEx OD316-5 for the Certified Service Facility Scheme

- Operational Document IECEx OD422 for the IECEx Conformity Mark Licensing System

IECEx ExCB/ExTL assessment report for KGS

INTERNATIONAL
ELECTROTECHNICAL
COMMISSION

CONTENTS

1     Assessment information	6

1.1	Type of Body covered by this assessment: <retain appropriate marks>	6

1.2	Type of assessment: <retain appropriate marks>	6

1.3	Details of body	6

1.3.1	Country	6

1.3.2	Name of body	6

1.3.3	Name and title of nominated principal contact	6

1.4	Assessment information	6

1.4.1	Members of the assessment team	6

1.4.2	Place(s) of assessment	6

1.4.3	Assessment date(s)	6

1.5	Application information and background information on the assessment	6

1.6	Scopes	7

1.6.1	ExCB scope for equipment certification scheme	7

1.6.2	ExTL scope	7

1.6.3	ExCB scope for Service Facilities Scheme	8

1.6.4	ExCB scope for ExMark Scheme	8

2	Common information	9

2.1	Legal entity of body	9

2.2	Financial support	9

2.3	History	9

2.4	Documentation	9

2.4.1	Quality manual	9

2.4.2	Procedures	9

2.4.3	Work instructions	9

2.4.4	Records (including test records where relevant)	9

2.4.5	Document change control	10

2.5	Confidentiality	10

2.6	Publications (Hard cover and Electronic)	10

2.7	Recognition and agreements	10

2.8	Internal audit and periodic management review	10

2.9	Contracting, subcontracting, use of other labs and use of other locations	10

2.10	Training and competence	10

2.11	Complaints and appeals (including appeals to IECEx)	11

2.12	Special facts to be noted	11

2.12.1	Scope extension for ISO 80079-36 and ISO 80079-37	11

2.12.2	Supporting documentation	11

2.13	Recommendations	11

3	ExCB for IECEx Certified Equipment Scheme	13

3.1	Assessment references	13

3.1.1	General references	13

3.1.2	Additional references applied for this assessment	13

3.2	ExCB persons interviewed	13

3.3	Associated ExTL(s)	13

3.4	Associated certification functions	13

3.5	National marks and certificates	13

There is a national mark and certificate for	13

3.6	Standards accepted	13

3.7	National differences to IEC standards	14

3.8	Organisation	14

3.8.1	Names, titles and experience of the senior executives	14

3.8.2	Name, title and experience of the quality management representative	14

3.8.3	Name and title of signatories for certification	14

3.8.4	Other employees in ExCB activity	14

3.9	Organizational structure	14

3.10	Administration	14

3.10.1	Administrative structure	14

3.10.2	Indemnity insurance	14

3.11	Resources	14

3.12	Committees (such as governing or advisory boards)	15

3.13	Certification operations	15

3.13.1	National approval/certification methods	15

3.13.2	Certification policy	15

3.13.3	Application for certification	15

3.13.4	Certification decision	15

3.13.5	Suspension and cancellation of certificates	15

3.14	Certificates issued	15

3.15	National accreditation	16

3.16	Assessment of manufacturers and issue of QARs	17

3.17	Comments (including issues found during assessment)	17

4	ExTL for IECEx Certified Equipment Scheme	18

4.1	Assessment references	18

4.1.1	General references	18

4.1.2	Additional references applied for this assessment	18

4.2	ExTL persons interviewed	18

4.3	Associated ExCB(s)	18

4.4	Organisation	18

4.4.1	Names, titles and experience of the senior executives	18

4.4.2	Name, title and experience of the quality management representative	18

4.4.3	Other employees in ExTL activity	19

4.5	Organizational structure	19

4.6	Resources	19

4.7	Test reports issued	19

4.8	National accreditation	20

4.9	Calibration	20

4.10	Uncertainty	20

4.11	Tests witnessed during the assessment visit	20

4.12	Participation in IECEx Proficiency Testing Programs	21

4.13	Comments (including issues found during assessment)	21

5	ExCB for Certified Service Facilities Scheme	22

5.1	Assessment references	22

5.2	ExCB persons interviewed	22

5.3	Organisation	22

5.3.1	Names, titles and experience of the senior executives	22

5.3.2	Name, title and experience of the quality management representative	22

5.3.3	Name and title of signatories for certification	22

5.3.4	Other employees in ExCB activity	22

5.4	Organizational Structure	23

5.5	Resources	23

5.6	Certification operations	23

5.6.1	National approval/certification Methods	23

5.6.2	Application for certification	23

5.6.3	Withdrawal and cancellation of certificates	23

5.7	Statistics	23

5.8	National accreditation	23

5.9	Comments (including issues found during assessment)	23

6	Annexes	24

Annex A Overall Organisation Chart	25

Annex B Organisation Chart of ExCB and ExTL	26

Annex C Accreditation Certificate for ISO/IEC 17065	27

Annex D Accreditation Certificate for ISO/IEC 17025	30

# Assessment information

## Type of Body covered by this assessment: <retain appropriate marks>

| ExCB for IECEx Certified Equipment Scheme |  |
| --- | --- |
| ExTL for IECEx Certified Equipment Scheme |  |
| ExCB for IECEx Certified Service Facilities Scheme |  |
| ExCB for IECEx Conformity Mark Licensing System |  |

NOTE 1	ExCB - IECEx Certification Body

NOTE 2	ExTL - IECEx Testing Laboratory

## Type of assessment: <retain appropriate marks>

| Pre-assessment for candidate body |  |
| --- | --- |
| Initial assessment for candidate body |  |
| Surveillance |  |
| Re-assessment |  |
| Scope extension |  |

## Details of body

### Country

Republic of Korea

### Name of body

Korea Gas Safety Corporation (KGS)

### Name and title of nominated principal contact

| Name | Title | E-mail address |
| --- | --- | --- |
| Gihoi Kim | General Manager | Kgh0803@kgs.or.kr |

## Assessment information

### Members of the assessment team

| Name | Role |
| --- | --- |
| Ron Webb | IECEx Lead Assessor |
| Gordana Ostojic | IECEx Assessor |

### Place(s) of assessment

1390 Wonjung-ro, Maengdong-myeon, Eumseong-gun, Chungcheongbuk-do, Rep of Korea

### Assessment date(s)

20-23 November 2017 and 17 April 2018 (Ron Webb)

## Application information and background information on the assessment

Re-assessment and Scope extension

## Scopes

### ExCB scope for equipment certification scheme

| Number | Title | Comments, eg if scope change |
| --- | --- | --- |
| IEC 60079-0 Edition 6 | Explosive atmospheres - Part 0: Equipment - General requirements |  |
| IEC 60079-1 Edition 7 | Explosive atmospheres - Part 1: Equipment protection by flameproof enclosures “d” |  |
| IEC 60079-2 Edition 6 | Explosive atmospheres - Part 2: Equipment protection by pressurized enclosure «p» |  |
| IEC 60079-6 Edition 4 | Explosive atmospheres - Part 6: Equipment protection by liquid immersion «o» |  |
| IEC 60079-7 Edition 5 | Explosive atmospheres - Part 7: Equipment protection by increased safety "e" |  |
| IEC 60079-11 Edition 6 | Explosive atmospheres - Part 11: Equipment protection by intrinsic safety “i” |  |
| IEC 60079-15 Edition 4 | Explosive atmospheres – Part 15: Equipment protection by type of protection "n" |  |
| IEC 60079-18 Edition 4 | Explosive atmospheres – Part 18: Equipment protection by encapsulation “m” |  |
| IEC 60079-26 Edition 3 | Explosive atmospheres - Part 26: Equipment with equipment protection level (EPL) Ga |  |
| IEC 60079-31 Edition 2 | Explosive atmospheres – Part 31: Equipment dust ignition protection by enclosure "t" |  |
| ISO 80079-36 Edition 1 | Explosive atmospheres – Part 36: Non-electrical Equipment for explosive atmospheres - Basic method and requirements | Scope Extension |
| ISO 80079-37 Edition 1 | Explosive atmospheres – Part 37: Non-electrical Equipment for explosive atmospheres – Non-electrical type of protection constructional safety “c”, control of ignition source “b”, liquid immersion “k” | Scope Extension |

NOTE 1	Unless otherwise indicated, earlier editions of standards (even if with a different number) are considered to be covered in the above scope for the purposes of the assessment.

NOTE 2	The above list highlights any extension of scope in the list above for new standards or later editions of standards already in scope.

### ExTL scope

The ExTL scope is the same as for the ExCB.

### ExCB scope for Service Facilities Scheme

| 60079-19 Edition 3 | Explosive atmospheres – Part 19: Equipment repair, overhaul and reclamation |
| --- | --- |
| Plus the following types of protection | ID |
| Flameproof Enclosure "d" | d |
| Increased Safety "e" | e |
| Type of Protection "n" | n |
| Pressurisation "p" | p |

### ExCB scope for ExMark Scheme

Not in scope

# Common information

## Legal entity of body

KGS is a public body established under Article 28 of the High Pressure Gas Safety Control Law.  There is a certificate of registration 133-82-02018 with the Taxation Department for KGS.

## Financial support

KGS receives about 40% financial support from the Government.  The rest is met from inspection, training, R&D, testing and certification income.

## History

Initially established in 1974 with the name KGS being used since 1979.

In 1995, the Institute of Gas Safety Technology was established as an annex research centre.

In 1998, it became an affiliate of the Ministry of Commerce, Industry and Energy, and was designated as a certification body for electrical explosion-proof performance.

In 2003, KGS was accepted in the IECEx Certified Equipment Scheme as an ExCB and an ExTL

In 2012, KGS was accepted in the IECEx Service Facility Scheme as an ExCB.

## Documentation

### Quality manual

KGS established a comprehensive quality management system to operate Ex certification and testing in accordance with ISO/IEC 17025 and ISO/IEC 17065 with each section numbered Ex-M-xx.

Checklists against ISO/IEC 17025 and ISO/IEC 17065 were prepared and found to meet the requirements of each section.

### Procedures

The next level in the system is the procedures which are numbered Ex-P-xx.

### Work instructions

The next level in the system is the instructions which are numbered Ex-I-xx.

### Records (including test records where relevant)

KGS has established documented procedure Ex-P-13 to specify the requirement on records.

All records related to Ex testing and certification activities are retained a minimum of 10 years in hard copy.

### Document change control

Document control is covered in Ex-P-04. All procedures and related documents are in printed form and subject to a document change control process. Document control was reviewed during the assessments and found to meet the requirements of the IECEx.

## Confidentiality

Confidentiality is addressed in Ex-M-27.  Each staff member is required to sign 	a confidentiality agreement.  The file containing the signed forms was viewed and found to be acceptable in meeting IECEx Requirements.

## Publications (Hard cover and Electronic)

There is a magazine that is circulated by KGS every month.

## Recognition and agreements

There are mutual recognition agreements in place with a number of bodies including

Baseefa, CSA (Sira) in the UK,

LCIE in France

NANIO CCVE in Russia

NEPSI in China

Simtars in Australia

FTZU in the Czech Republic

PTB, Dekra Exam in Germany

Dekra in the Netherlands

INTERTEK in US

CSA in Canada

IMQ in Italy.

## Internal audit and periodic management review

The Internal audit is in Ex-M-14 and Ex-P-14.

The internal audit for 13~15 Feb. 2017 was viewed.  It used a lead ISO auditor and a technical expert who is also an external auditor of manufacturers (not in the Ex field).  The audit indicates it covers ISO/IEC 17025 and 17020.  NCRs were found but not in the Ex department.

The Management review is in Ex-M-15 and Ex-P-15 and which is carried out once a year. The review carried out on 28 April 2017 was viewed and found to meet the requirements of the IECEx.

## Contracting, subcontracting, use of other labs and use of other locations

KGS subcontract with KTL for the measurement of screw threads, vibration tests for bi-pin tubes in luminaires, the impact test for Ex e batteries and the charging test in IEC 60079-0. There is no sub-contracting to other bodies.  There is a file covering the subcontracting to KTL which shows the tests covered, the relevant accreditation from KOLAS is included in file and there is a signed contract that also includes a confidentiality agreement.

Witness testing requirements are contained in Ex-P-18 Certification and Test.
It was found to use the methods required by OD24.

## Training and competence

Training and competence is addressed in Ex-M-17 and Ex-P-16.

The computer system is used to record all external training, duty statements and qualifications.

A plan is developed for internal and external training each year and recorded in a file.  The file shows the information about who attended the training and the information about what was covered.

Instruction Ex-I-04 Operating audit addresses the assigning of competency to staff members. Instruction Ex-I-06 Certification of Employee addresses the assigning of competency to staff members.

Ex-P-23 Assuring the Quality also covers how a plan is made for employees

## Complaints and appeals (including appeals to IECEx)

Procedure Ex-P-09 addresses appeals, complaints and disputes.  It includes specific reference to the potential use of the IECEx Board of Appeal by the ExCB or the applicant. The process is managed by the Quality Manager.  There have not been any Ex related complaints.

## Special facts to be noted

### Scope extension for ISO 80079-36 and ISO 80079-37

Interviews with staff members regarding non-electrical equipment showed that they had a good understanding of the concept and a need an Ignition Hazard assessment to be prepared, usually by the manufacturer in conjunction with KGS.

Although no certificates had been prepared, KGS staff carried out an Ignition Hazard Assessment on an electrically driven pressure pump.  The result was reviewed by the assessment team and found to be acceptable.  They also prepared a report on the pump using the IECEx templates.  Again, this was seen to be acceptable.

It was noted that the requirements of OD 280 had been incorporated into the quality system.

### Supporting documentation

Copies of additional supporting information for this assessment have been provided to the applicant and the IECEx Secretariat.  These are included in a site assessment report and include:

- Details of issues raised and how these have been resolved

- Checklist for ISO/IEC 17065

- Checklist for ISO/IEC 17025

- Completed TCD

- Photos of the facilities/tests witnessed are included in the TCD

- Assessors’ notes

## Recommendations

Based on the assessments performed on 20-23 November 2017 and a further visit on 17 April 2018, which was made to carry out to finalise outstanding issues as well as carry out a further scope extension, KGS is recommended for continued acceptance in the IECEx scheme as:

- An ExCB in the IECEx Certified Equipment Scheme

- An ExTL in the IECEx Certified Equipment Scheme

- An ExCB in the IECEx Certified Service Facilities Scheme

This is according to the scope of the standards listed in this document including the extension of scope.

As the scope extension standards are not yet on the KOLAS accreditation KGS will be subject to annual surveillance by an IECEx assessment team. This first Surveillance visit to be carried out within six months.

| Ron Webb | Gordana Ostojic |
| --- | --- |
| Lead Assessor | Expert Assessor |

Date:  17 April 2018

# ExCB for IECEx Certified Equipment Scheme

## Assessment references

### General references

- IECEx02 IECEx Certified Equipment Scheme covering equipment for use in explosive atmospheres – Rules of Procedure

- OD003-2  Assessment, surveillance assessment and re-assessment of ExCBs and ExTLs operating in the IECEx 02, IECEx Certified Equipment Scheme

- ISO/IEC 80079-34 Edition 1, Explosive atmospheres – Part 34: Application of quality systems for equipment manufacture

- OD009 Issuing of CoCs, ExTRs and QARs

- IECEx Document OD 025 Guidelines on the Management of Assessment and Surveillance programs for the assessment of Manufacturer’s Quality Systems in accordance with the IECEx Scheme

- OD0026 IECEx Certified Equipment Scheme – Guidelines for the qualification of Lead Auditor and Auditors, in accordance with the IECEx System

- ISO/IEC 17065: 2012, Edition 1, General requirements for bodies operating product certification systems Conformity assessment — Requirements for bodies certifying products, processes and services

- IECEx Technical Capability Document (TCD)

- ExTAG decision sheets (DSs)

NOTE	The latest editions of the above documents were applied

### Additional references applied for this assessment

- OD 280:2016 IECEx Certified Equipment Scheme – Guide to Certification of Non-Electrical Equipment and Protective Systems

## ExCB persons interviewed

| Name | Position |
| --- | --- |
| Gihoi Kim | Department Leader, Quality Manger |
| Inchoon Park | Manger |

## Associated ExTL(s)

The KGS ExTL is integral with the ExCB

## Associated certification functions

KGS issues certificates for gas appliances; it issues certification to ISO 9001 for all areas; to ISO 14000; and national Ex certification.

## National marks and certificates

## There is a national mark and certificate for

## Standards accepted

See clause 1.6 of this report

## National differences to IEC standards

National differences to IEC standards are those listed in the latest version of the IECEx Scheme Bulletin.

## Organisation

### Names, titles and experience of the senior executives

| Name | Title | Experience |
| --- | --- | --- |
| JongSam Moon | Director | 32 years at KGS |
| Gihoi Kim | Department Leader | 22 years at KGS and 10 years in Ex |

### Name, title and experience of the quality management representative

| Name | Title | Experience |
| --- | --- | --- |
| Gihoi Kim | General Manager, Ex Testing and Certification Department and Quality Manager | 22 years at KGS and 10 years in Ex |

### Name and title of signatories for certification

| Name | Title | Comments |
| --- | --- | --- |
| Gihoi Kim | As above | Kgh0803@kgs.or.kr |

### Other employees in ExCB activity

| Name | Title | Responsibility and Experience in Ex |
| --- | --- | --- |
| Inchoon Park | Manger | 7 years at KGS and 6 years in Ex |

## Organizational structure

The total staff of the ‘Hazardous Locations’ group in KGS is now 11 with two personnel in the ExCB.  The most recent change in structure was in January 2015 when the reporting structure of the group was changed and Gihoi Kim became Department Leader of the group in addition to his role as head of the ExCB.

The organisational structures are shown in Annex A and B of this report and also appear in Ex-M-02 of the Manual.

## Administration

### Administrative structure

See Annex B

### Indemnity insurance

KGS indemnity is covered by the government as it is a government agency.  In addition it has an indemnity insurance cover for of KRW10,000,000,000.  This policy was viewed.

## Resources

Principally there are 2 dedicated persons in the ExCB plus additional technical resources used as KGS Audit Resources plus dedicated admin support. These resources cover both the IECEx 02 and 03 Schemes.

## Committees (such as governing or advisory boards)

There is an internal audit group for auditing of finances and implementation of policies, compliance with regulations etc.  The head of this group is appointed by the Government.

KGS also has a board that covers a range of high level functions, including supervision of finances, approval of major capital expenditure, major policy issues and the structure of KGS.  It is not directly involved with certification functions.

## Certification operations

### National approval/certification methods

There is a national mark and certificate for Ex

### Certification policy

The quality policy is defined in the Manual in Ex-M-0B.  The policy includes reference to ISO/IEC 17025 and 17065 and specific reference to certification and testing.  The policy is signed by the General Manager under delegation from the President.

### Application for certification

KGS has an application form for clients to use when applying for IECEx certification. The procedures for dealing with applications for certification are contained in Ex-P-05 Application and contract review.  This document defines the responsibilities of the relevant personnel and the scope.

Feedback is obtained in accordance with procedure Ex-P-08 Service to the Client.  This is done by means of a questionnaire.

As seen in Annex B, separation between ExCB and ExTL is achieved through the use of different personnel in each area.  In addition, as noted earlier, the certification committee also provides an independent role in the certification process.

### Certification decision

There is a Certification Committee with its terms of reference defined in the procedure Ex-P-18 Certification and Test (5.5, 6.5).  The General Manager is the Chairman of this committee.  It comprises minimum 5 members, including the General Manager from KGS and 4 other people both from within and outside the Ex Testing & Certification Service Department.  There are no external people.

### Suspension and cancellation of certificates

Ex-I-02 Maintenance of Certification Registration addresses suspension, withdrawal and cancellation of certificates and ExTRs.  The decision on cancellation is taken by the Certification Committee

## Certificates issued

Number of certificates issued under for the preceding four years for each type of protection.

| Standard numbers | Type of protection or other identifying information | Number of issued certificates (for last 4 years) |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
|  |  | 2014 | 2015 | 2016 | 2017 | Total |
| IEC 60079-0 | Explosive atmospheres – Part 0 : Equipment – General requirements | 20 | 27 | 40 | 19 | 106 |
| IEC 60079-1 | Explosive atmospheres – Part 1 : Equipment protection by flameproof enclosures “d” | 18 | 24 | 33 | 13 | 88 |
| IEC 60079-2 | Explosive atmospheres – Part 2 : Equipment protection by pressurized enclosure “p” | 0 | 0 | 0 | 0 | 0 |
| IEC 60079-6 | Explosive atmospheres – Part 6 : Equipment protection by oil immersion “o” | 0 | 0 | 0 | 0 |  |
| IEC 60079-7 | Explosive atmospheres – Part 7 : Equipment protection by increased safety “e” | 12 | 20 | 24 | 11 | 67 |
| IEC 60079-11 | Explosive atmospheres – Part 11 : Equipment protection by intrinsic safety “i” | 0 | 0 | 0 | 1 | 1 |
| IEC 60079-15 | Explosive atmospheres – Part 15 : Equipment protection by type of protection “n” | 0 | 1 | 0 | 3 | 4 |
| IEC 60079-18 | Explosive atmospheres – Part 18 : Equipment protection by encapsulation “m” | 8 | 11 | 20 | 4 | 43 |
| IEC 60079-26 | Explosive atmospheres – Part 26 : Equipment with protection level (EPL) Ga | 0 | 0 | 0 | 0 | 0 |
| IEC 60079-31 | Explosive atmospheres – Part 31 : Equipment dust ignition protection by enclosure “t” | 2 | 4 | 6 | 4 | 16 |

NOTE	Above include certificates to IEC 60079-0 unless otherwise shown

National Certificates issued

|  | 2014 | 2015 | 2016 | Oct 31 2017 | Total |
| --- | --- | --- | --- | --- | --- |
| Ex d | 239 | 428 | 454 | 207 | 1328 |
| Ex p | 31 | 105 | 78 | 76 | 290 |
| Ex e | 137 | 152 | 179 | 112 | 580 |
| Ex i | 0 | 0 | 8 | 5 | 13 |
| Ex n | 33 | 55 | 43 | 52 | 183 |
| Ex m | 0 | 0 | 0 | 5 | 5 |
| Ex t | 88 | 153 | 119 | 76 | 436 |

## National accreditation

The Government authorises KGS to issue National Ex certificates.  They also have formal accreditation as a certifying body to ISO/IEC 17065 from the Korea Accreditation System KAS

– See Annex C.

This is valid until 20 April 2021.  The scope of accreditation that forms part of the certificate does not explicitly mention Ex activities although section 289 is stated to be relevant for Ex equipment. The scope Category covers sections:

281 – Electrical equipment manufacturing: motor, generator, electrical converter, electrical supply controller manufacturing.

284 – Bulb or lighting fixtures manufacturing: Lighting fixture manufacturing

289 – Other electrical equipment manufacturing: Other electrical equipment manufacturing.

Mechanical (non-electrical) equipment is not mentioned. Therefore, on this basis an annual surveillance assessment is to be carried out.

## Assessment of manufacturers and issue of QARs

The manufacturing assessments are generally done by ExCB staff, but use may be made by Ex experts from the ExTL. Staff are also under training as an IECEx assessor of manufacturers.  The procedure for this is Ex-I-01 and 04.

## Comments (including issues found during assessment)

There were some issues found which were cleared to the satisfaction of the assessment team. These included documentation regarding OD024 and subcontract agreements.

# ExTL for IECEx Certified Equipment Scheme

## Assessment references

### General references

- IECEx02 IECEx Certified Equipment Scheme covering equipment for use in explosive atmospheres – Rules of Procedure

- IECEx OD003-2 Assessment, surveillance assessment and re-assessment of ExCBs and ExTLs operating in the IECEx 02, IECEx Certified Equipment Scheme

- IECEx OD009 Issuing of CoCs, ExTRs and QARs

- ISO/IEC 17025:2005 Edition 2, General requirements for the competence of testing and calibration laboratories

- IECEx Technical Capability Document (TCD)

- ExTAG decision sheets (DSs)

- OD 202 IECEx Certified Equipment Scheme – IECEx Proficiency Testing Program

NOTE	The latest editions of the above documents were applied.

### Additional references applied for this assessment

a)	OD 280:2016 IECEx Certified Equipment Scheme – Guide to Certification of Non-Electrical Equipment and Protective Systems

## ExTL persons interviewed

| Name | Position |
| --- | --- |
| Myonju Choi | Senior Manger |
| Woosuk Kim | Manager |
| Jonghyun Kim | Assistant Manager |
| Heejeong Choi | Assistant Manager |
| Sooyoung Kwon | Trainee |
| Seonghun Shim | Assistant Manager |

## Associated ExCB(s)

The KGS ExCB is integral with the ExTL

## Organisation

### Names, titles and experience of the senior executives

| Name | Title | Experience |
| --- | --- | --- |
| JongSam Moon | Director | 32 years at KGS |
| Gihoi Kim | Department Leader | 22 years at KGS and 10 years in Ex |

### Name, title and experience of the quality management representative

| Name | Title | Experience |
| --- | --- | --- |
| Gihoi Kim | General Manager, Ex Testing and Certification Department and Quality Manager | 22 years at KGS and 10 years in Ex |

### Other employees in ExTL activity

| Name | Title/responsibility | Experience in Ex |
| --- | --- | --- |
| Changhun Yoo | Technical Manager of ExTL | 22 years at KGS and 9 years in Ex |

## Organizational structure

See Annex B

## Resources

The ExTL has a shared General Manager with the ExCB and eight staff.  It is well resourced with test equipment.

## Test reports issued

Number of test reports (ExTRs) issued under for the preceding four years for each type of protection.

| Standard numbers | Type of protection or other identifying information | Number of issued reports (ExTRs) (for last 4 years) |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
|  |  | 2014 | 2015 | 2016 | 2017 | Total |
| IEC 60079-0 | Explosive atmospheres – Part 0 : Equipment – General requirements | 20 | 27 | 40 | 19 | 106 |
| IEC 60079-1 | Explosive atmospheres – Part 1 : Equipment protection by flameproof enclosures “d” | 18 | 24 | 33 | 13 | 88 |
| IEC 60079-2 | Explosive atmospheres – Part 2 : Equipment protection by pressurized enclosure “p” | 0 | 0 | 0 | 0 | 0 |
| IEC 60079-6 | Explosive atmospheres – Part 6 : Equipment protection by oil immersion “o” | 0 | 0 | 0 | 0 | 0 |
| IEC 60079-7 | Explosive atmospheres – Part 7 : Equipment protection by increased safety “e” | 12 | 20 | 24 | 11 | 67 |
| IEC 60079-11 | Explosive atmospheres – Part 11 : Equipment protection by intrinsic safety “i” | 0 | 0 | 0 | 1 | 1 |
| IEC 60079-15 | Explosive atmospheres – Part 15 : Equipment protection by type of protection “n” | 0 | 1 | 0 | 3 | 4 |
| IEC 60079-18 | Explosive atmospheres – Part 18 : Equipment protection by encapsulation “m” | 8 | 11 | 20 | 4 | 43 |
| IEC 60079-26 | Explosive atmospheres – Part 26 : Equipment with protection level (EPL) Ga | 0 | 0 | 0 | 0 | 0 |
| IEC 60079-31 | Explosive atmospheres – Part 31 : Equipment dust ignition protection by enclosure “t” | 2 | 4 | 6 | 4 | 16 |

NOTE 1	Above include reports to IEC 60079-0 unless otherwise shown

Local Reports - only internal use. These test reports are not issued to the clients:

|  | 2014 | 2015 | 2016 | 2017 | Total |
| --- | --- | --- | --- | --- | --- |
| Ex d | 239 | 428 | 454 | 207 | 1328 |
| Ex p | 31 | 105 | 78 | 76 | 290 |
| Ex e | 137 | 152 | 179 | 112 | 580 |
| Ex i | 0 | 0 | 8 | 5 | 13 |
| Ex n | 33 | 55 | 43 | 52 | 183 |
| Ex m | 0 | 0 | 0 | 5 | 5 |
| Ex t | 88 | 153 | 119 | 76 | 436 |

## National accreditation

KGS has national accreditation according ISO/IEC 17025 from KOLAS number KT186 (see Annex D).  At the time of the assessment visit the scope of this accreditation did not cover the latest editions. The assessment was carried out in Oct 2017.Mechanical (non-electrical) equipment is not mentioned. Therefore, on this basis an annual surveillance assessment is to be carried out.

## Calibration

All equipment used to take measurements is calibrated.  Calibration is controlled by procedure EX-P-20.  The certification is carried out by external laboratories having National Accreditation. Sample certificates were viewed and found to be in accordance with the requirements.

## Uncertainty

Procedure EX-P-19 covers uncertainty.   It refers to OD012. Sample uncertainty budgets were viewed.

## Tests witnessed during the assessment visit

The following tests were witnessed during the assessment visit:

| Standard and edition | Clause number | Test | Comments |
| --- | --- | --- | --- |
| IEC 60079-1 Ed 7 | 15.2.2.2 | Reference pressure determination | IIC test carried out successfully |
| IEC 60079-1 Ed 7 | 15.3.3.2 | Non-transmission test | IIC test carried out successfully |
| IEC 60079-0 Ed 6 | 26.13 | Surface resistance | Test carried out successfully |
| IEC 60079-18 Ed 4 | 8.1.2 | Dielectric strength Ex m | Test carried out successfully |
| IEC 60079-0 Ed 6 | 26.4.5.1 | Dust test IP5X/6X | Test carried out successfully |
| IEC 60079-0 Ed 6 | 26.5.1 | Temperature rise Ex d lamp | Set up in draught proof room. Survey with IR to determine hotspots |
| IEC 60079-11 Ed 6 | 10.1.3.2 | STA calibration for IIC, 1.5 SF | Calibration successful. Ignition at 240 revolutions |

## Participation in IECEx Proficiency Testing Programs

Program: PTB Ex PT Scheme

| Year(s) of participation | IECEx Proficiency Testing program | General information about results |
| --- | --- | --- |
| 2010 | Program “Explosion Pressure – Test Round 2010” | Finished - Acceptable |
| 2010 | Program “Spark Ignition – Test Round 2010” | Finished - Acceptable |
| 2013 | Program “Flame Transmission – Test Round 2013” | Finished - Acceptable |
| 2013 | Program “Temperature Classification – Test Round 2013” | Finished – Acceptable following discussions with PT provider |
| Not participated | Program “Electrostatic Charge – Test Round 2015” |  |
| 2015 | Program “Intrinsic Safety – Test Round 2015” | Finished - Acceptable |
| Registered | Program “Explosion Pressure – Test Round 2017” * |  |
| Not registered | Program “Pressurized Enclosure – Test Round 2017” * |  |

## Comments (including issues found during assessment)

There were some issues that were cleared to the satisfaction of the assessment team.

These included:

the need for clear records of test conditions and measurements

the need to give clarity to test procedures

the provision of full justification where tests are waived

# ExCB for Certified Service Facilities Scheme

## Assessment references

- IECEx 03-5  IECEx Certified Service Facilities Scheme covering repair and overhaul of Ex equipment – Rules of Procedure

- IECEx OD 313-5, IECEx Certified Service Facilities Scheme – Part 5: Repair, overhaul and reclamation of Ex equipment, Assessment and Certification of Service Facilities involved in the repair, overhaul and reclamation of Ex equipment – Procedures

- IECEx OD 314-5, IECEx Certified Service Facilities Scheme – Part 5: Repair, overhaul and reclamation of Ex equipment, Quality Management System requirements for IECEx Service Facilities involved in the repair, overhaul and reclamation of Ex equipment IECEx

- IECEx OD 315-5, IECEx Certified Service Facilities Scheme – Part 5: Repair, overhaul and reclamation of Ex equipment, Additional requirements for IECEx Service Facilities involved in the repair, overhaul and reclamation of Ex equipment

- IECEx OD 316-5, IECEx Certified Service Facilities Scheme – Part 5: Repair, overhaul and reclamation of Ex equipment, Assessment procedures for IECEx acceptance of Candidate Certification Bodies (ExCBs) for the purpose of issuing IECEx Certificates to Ex Service Facilities involved in the repair, overhaul and reclamation of Ex equipment

- ISO/IEC 17065, General requirements for bodies operating product certification systems

- IEC 60079-19 Explosive atmospheres – Part 19: Equipment repair, overhaul and reclamation

NOTE	The latest editions of the above documents were applied

## ExCB persons interviewed

| Name | Position |
| --- | --- |
| Myonju Choi | Senior Manager |
| Woosok Kim | Manager |

## Organisation

### Names, titles and experience of the senior executives

| Name | Title | Experience |
| --- | --- | --- |
| JongSam Moon | Director | 32 years at KGS |
| Gihoi Kim | Department Leader | 22 years at KGS and 10 years in Ex |

### Name, title and experience of the quality management representative

| Name | Title | Experience |
| --- | --- | --- |
| Gihoi Kim | General Manager, Ex Testing and Certification Department and Quality Manager | 22 years at KGS and 10 years in Ex |

### Name and title of signatories for certification

| Name | Title | Comments |
| --- | --- | --- |
| Gihoi Kim | As above | Kgh0803@kgs.or.kr |

### Other employees in ExCB activity

| Name | Title/responsibility | Experience in Ex |
| --- | --- | --- |
| Myonju Choi | Senior Manager | 22 years at KGS and 6 years in Ex |
| Woosok Kim | Manager | 7 years at KGS and 7 years in Ex |

## Organizational Structure

See Annex B

## Resources

See 3.11

## Certification operations

### National approval/certification Methods

N/A

### Application for certification

KGS has an application form for clients to use when applying for IECEx certification. The procedures for dealing with applications for certification are contained in Ex-P-05 Application and contract review.  This document defines the responsibilities of the relevant personnel and the scope.

### Withdrawal and cancellation of certificates

Ex-I-02 Maintenance of Certification Registration addresses withdrawal and cancellation of certificates.  The decision on cancellation is taken by the Certification Committee

## Statistics

KGS have carried out five facilities assessment and issued a certificate during past 2 years for the types of protection shown below

flameproof		d

increased safety	e

type of protection	n

pressurization		p

## National accreditation

N/A

## Comments (including issues found during assessment)

none

# Annexes

See Contents.  (add,modify or delete annexes as necessary).  Please note the following

- Overall Organisation Chart

- Organisation Chart of ExCB and ExTL

- Accreditation Certificate for ISO/IEC 17065

- Accreditation Certificate for ISO/IEC 17025