INTERNATIONAL ELECTROTECHNICAL COMMISSION SYSTEM FOR CERTIFICATION TO STANDARDS RELATING TO EQUIPMENT FOR USE IN EXPLOSIVE ATMOSPHERES (IECEx SYSTEM)

TITLE: Re-assessment Report for the continued acceptance of KSC Co., Ltd an Accepted ExTL within the IECEx Equipment Scheme 02.

Circulation to: Members of the IECEx Management Committee, ExMC

INTRODUCTION

In accordance with the 5 Year re-assessment plan for the surveillance and monitoring of bodies within the IECEx System, the following document contains the IECEx Reassessment Report for the continued acceptance KSC Co., Ltd an Accepted ExTL within the IECEx Equipment Scheme 02.

This report is hereby submitted for endorsement during the 2023 ExMC Meeting.

Chris Agius

IECEx Secretariat

| IECEx Secretariat Australia Square Level 33, 264 George Street Sydney. NSW 2000 Australia | Tel: +61 2 4628 4690 Fax: +61 2 4625 3480 Email: info@iecex.com |
| --- | --- |

IEC System for certification to standards relating to equipment for use in Explosive Atmospheres (IECEx System)

IECEx Assessment Report

IECEx assessment report form for use by IECEx assessment teams to report assessments conducted according to the relevant IECEx assessment procedures of:

Operational Document IECEx OD 003-2 for the Certified Equipment Scheme

Operational Document IECEx OD 316-* for the Certified Service Facility Scheme

Operational Document IECEx OD 422 for the IECEx Conformity Mark Licensing Scheme

Operational Document IECEx OD 501 for the Personnel Competence Scheme

IECEx ExTL assessment report for

KSC Co., Ltd

1106, Cheoinseong-ro, Namsa-myeon, Cheoin-gu,

Yongin-si, Gyeonggi-do,17115,

Republic of Korea

CONTENTS

1	Assessment information	5

1.1	Type of body covered by this assessment:	5

1.2	Type of assessment:	5

1.3	Details of body	5

1.3.1	Country	5

1.3.2	Name of body	5

1.3.3	Name and title of nominated principal contact	5

1.4	Assessment information	5

1.4.1	Members of the assessment team	5

1.4.2	Place(s) of assessment	5

1.4.3	Assessment date(s)	6

1.5	Application information and background information on the assessment	6

1.6	Scopes	6

1.6.1	ExCB scope for equipment certification scheme	6

1.6.2	ExTL scope	6

1.6.3	ATF Scope – N/A	6

1.6.4	ExCB scope for Service Facilities Scheme – N/A	6

1.7	ExCB scope for Conformity Mark Licensing Scheme – N/A	6

1.8	ExCB scope for IECEx Personnel Competence Scheme – N/A	6

2	Common information	7

2.1	Legal entity of body	7

2.2	Financial support	7

2.3	History	7

2.4	Documentation	7

2.4.1	Quality manual	7

2.4.2	Procedures	8

2.4.3	Work instructions	8

2.4.4	Records (including test records where relevant)	8

2.4.5	Document change control	8

2.5	Confidentiality	8

2.6	Communication with public and customers (Hard copy and Electronic)	8

2.7	Recognitions and agreements	9

2.8	Internal audit	9

2.9	Management review	9

2.10	Contracting, subcontracting and witness testing	9

2.10.1	Contracting	9

2.10.2	Subcontracting	9

2.10.3	Off-site and Witness testing	9

2.11	Training and competence	9

2.12	Complaints and appeals (including appeals to IECEx)	10

2.13	Impartiality	10

2.14	Active involvement in development of Decision Sheets	10

2.15	Special facts to be noted	10

2.16	Supporting documentation	10

2.17	Recommendations	10

3	ExCB for IECEx Certified Equipment Scheme – N/A	11

4	ExTL for IECEx Certified Equipment Scheme	11

4.1	Assessment references	11

4.1.1	General references	11

4.1.2	Additional references applied for this assessment	11

4.2	ExTL persons interviewed	11

4.3	Associated ExCB(s)	11

4.4	Organisation	11

4.4.1	Names, titles and experience of the senior executives	11

4.4.2	Name, title and experience of the quality management representative	12

4.4.3	Other employees in ExTL activity	12

4.5	Organizational structure	12

4.6	Resources	12

4.7	Test reports issued	12

4.8	National accreditation	13

4.9	Calibration	13

4.10	Tests witnessed during the assessment visit	13

4.11	Participation in IECEx Proficiency Testing Programs	13

4.12	Comments (including issues found during assessment)	14

6	ExCB for Certified Service Facilities Scheme – N/A	15

7	IECEx Conformity Mark Licensing Scheme – N/A	15

8	Annexes	15

Annex A Scope for IECEx Certified Equipment Scheme	16

A.1	Current standards	16

A.2	Superseded standards	17

Annex B Overall Organisation Chart	18

Annex C Accreditation Certificate for ISO/IEC 17025	19

# Assessment information

## Type of body covered by this assessment:

| ExCB for IECEx Certified Equipment Scheme |  |
| --- | --- |
| ExTL for IECEx Certified Equipment Scheme |  |
| ATF for IECEx Certified Equipment Scheme |  |
| ExCB for IECEx Certified Service Facilities Scheme |  |
| ExCB for IECEx Conformity Mark Licensing System |  |
| ExCB for IECEx Certification of Personnel Competency Scheme |  |

NOTE 1	ExCB - IECEx Certification Body

NOTE 2	ExTL - IECEx Testing Laboratory

NOTE 3  ATF - Additional Testing Facility

## Type of assessment:

| Pre-assessment for candidate body |  |
| --- | --- |
| Initial assessment for candidate body |  |
| Surveillance |  |
| Re-assessment |  |
| Scope extension |  |

## Details of body

### Country

Korea, Republic of

### Name of body

KSC Co., Ltd

1106, Cheoinseong-ro, Namsa-myeon, Cheoin-gu,

Yongin-si, Gyeonggi-do,17115

### Name and title of nominated principal contact

| Name | Title | E-mail address |
| --- | --- | --- |
| Sanghee Kim | Chief Executive Officer | sh.kim@exksc.com |

## Assessment information

### Members of the assessment team

| Name | Role |
| --- | --- |
| Gordana Ostojic | IECEx Lead Assessor |

### Place(s) of assessment

Due to the COVID-19 pandemic related restrictions, the assessment is remotely carried out by using Microsoft Teams, IECEx OD-060 is applied for this assessment.

| 1106, Cheoinseong-ro, Namsa-myeon, Cheoin-gu, Yongin-si, Gyeonggi-do,17115, Republic of Korea | Remote audit |
| --- | --- |

### Assessment date(s)

Remote audit:

Day 1: 21 July 2022 ( 9:30 am to 12:30 pm;  and 1:30 pm- 4:30 pm; Sydney time)

Day 2: 22 July 2022 ( 9:30 am to 12:30 pm;  and 1:30 pm- 4:30 pm; Sydney time)

Day 3: 26 July 2022 ( 9:30 am to 12:30 pm;  and 1:30 pm- 4:30 pm; Sydney time)

Day 4: 27 July 2022 ( 9:30 am to 12:30 pm;  and 1:30 pm- 4:30 pm Sydney time)

Day 5: 29 July for closing meeting ( 9:30 am  – 12:30 pm Sydney time)

## Application information and background information on the assessment

Previous reports for the First assessment and for the Scope extension were provided by IECEx Secretariat. Due to the COVID-19 pandemic related restrictions, this re- assessment was remotely carried out by using Microsoft Teams, IECEx OD-060 is applied for this re-assessment.

## Scopes

### ExCB scope for equipment certification scheme

The following ExCBs are associated to this ExTL:

- KSC Poland

Chorzowska 150 Street,
40-101 Katowice
POLAND

- CML

Unit 1, Newport Business Park
New Port Road
Ellesmere Port
CH65 4LZ
UNITED KINGDOM

The scope for CML and PSC Poland as shown on www.iecex.com, covers the full scope of ExTL.

### ExTL scope

This ExTL is not integral with any of the above ExCB, however the ExTL works with each ExCB based on the associated cooperation agreement.

The scope of the ExTL is included in Annex A.

### ATF Scope – N/A

### ExCB scope for Service Facilities Scheme – N/A

## ExCB scope for Conformity Mark Licensing Scheme – N/A

## ExCB scope for IECEx Personnel Competence Scheme – N/A

# Common information

## Legal entity of body

KSC is Limited company, set up in 2010.

## Financial support

Income is solely from testing.

## History

In November 2010 Korea Safety Certification Company was established as a privately owned company, which aided foreign manufacturers in the preparation of documents for Korean Ex certificates.

In December 2013 KSC Co., Ltd was established. New laboratory was built dedicated for Testing in 2016. KSC established Quality Laboratory System in accordance with ISO 17025 and achieved KOLAS accreditation.

In July 2017, KSC has been accredited as IEC ExTL in IECEx equipment scheme.

## Documentation

### Quality manual

KSC has a comprehensive Quality manual KSC-KM-** that consists of the following sections:

| Doc. No. | Doc. Name | Contents |
| --- | --- | --- |
| KSC-KM-01 | General Requirements | I. General II. Impartiality III. Confidentiality IV. Statement of quality policy and impartiality |
| KSC-KM-02 | Structural requirements | I. Structural requirements |
| KSC-KM-03 | Resource requirements | I. General II. Personnel III. Facilities and environmental conditions IV. Equipment V. Metrological traceability VI. Externally provided products and services |
| KSC-KM-04 | Process requirements | I. Review of requests, tenders and contracts II. Selections, verification and validation of methods III. Sampling IV. Handling of test or calibration items V. Technical records VI. Evaluation of measurement uncertainty VII. Ensuring the validity of results VIII. Reporting of results IX. Complaints X. Nonconforming work XI. Control of data and information management |
| KSC-KM-05 | Management system requirements | I. Options II. Management system documentations III. Control of management system documents IV. Control of records V. Actions to address risks and opportunities VI. Improvement VII. Corrective actions VIII. Internal audits IX. Management reviews |

It addresses IECEx certified equipment scheme procedure rules to IECEx02.

### Procedures

Series of Operating procedures KSC-KP-** cover in addition to IECEx procedure Facility & Environment, Personnel, Internal audit, management review, Document control, Record control, Sample identification and handling, Calibration procedure. The procedures meet the requirements of the IECEx Scheme 02.

### Work instructions

The following work instructions related to Explosion protections were viewed:

KSC-KIT-39 – Reference pressure

KSC-KIT-10- Overpressure test

KSC-KIT-33-Test for non-transmission

KSC-KIT-01- Temperature measurement

KSC-KIT-38 – Spark test

KSC-KIT-83- Restricted breathing test

The Work instructions are appropriate and meet the requirements of the IECEx Scheme 02.

### Records (including test records where relevant)

The retention of records requirements are provided in procedure KSC-KM-04, where it is written 10 years from the time of product termination.

The records are stored in Hard (paper) documents and as well electronic media.

The system meets the process requirements of OD 207. The overall system meets the requirements of IECEx system.

### Document change control

Procedure KSC-KP-18 covers document change control. Procedure KSC-KPA-01 clause 4.18 covers how to keep up to date with changes in IECEx documents e.g. ODs, new ExTAG decision sheets and others.

The system meets the requirements of IECEx.

## Confidentiality

The procedure for human resources is provided in document KSC-KP-01. All laboratory staff and members from external resource make a documented pledge to ensure confidentiality according to procedure KSC-KP-01, the pledges are documented in form KSC-KP-01-F10. Records of signed Confidential agreement for staff were sighted and found to meet the requirements of the IECEx.

## Communication with public and customers (Hard copy and Electronic)

Website(exksc.com) can make a communication with public, sometimes participate in exhibition with hard copy pamphlet, and communicate with customers through electronic email.

## Recognitions and agreements

KSC has accreditation by A2LA in accordance with ISO/IEC 17025.

There are two agreements available:

- KSC as ExTL and KSCP as Associated CB

- KSC as ExTL and CML as Associated CB

The agreements between Certification Management Limited, CML, GB, and Korea Safety Certification Co., Ltd (KSC) were sighted and were found to meet IECEx requirements.

## Internal audit

Procedure KSC-KP-21 determines that an audit is to be carried out once a year. The last internal audit was carried out on 20 July 2021. There were two findings, and they were satisfactorily closed with corrective actions.

The system meets the requirements of ISO/IEC 17065, ISO/IEC 17025 and IECEx.

## Management review

Procedure is KSC-KP-22 specifies that the management review which is carried out at least once a year. The last management review was held on 10 January 2022.The minutes of this meeting include internal audits discussion, management review and evaluation of inputs and outputs, customer satisfaction.

The system meets the requirements of ISO/IEC 17065, ISO/IEC 17025 and IECEx.

## Contracting, subcontracting and witness testing

### Contracting

The following procedures addressing contracting are: KSC-KM-03 -Quality manual for Resource requirements, KSC-KP-07 which is operating procedure for externally provided products and services, KSC-KP-08 which is operating procedure for review of request, tenders and contracts, and KSC-KPA-01, which is operating procedure for IECEx work.

If needed, in the field of explosion-proof testing, the requirements for testing in non-fixed locations IECEx OD024 procedure is followed.

### Subcontracting

No subcontracting in the field of explosion protection testing as all tests are conducted by ExTL.

### Off-site and Witness testing

The procedure KSC-KPA-01 addresses the requirements for the Off-site and Witness testing, including the requirements and signed agreement by the IECEx OD 024 and assessment by the OD018. All Off-site testing facilities are registered in IECEx OD024 Testing register on-line.

## Training and competence

This is controlled according to the quality manual KSC-KM-03 (Resource requirements) and the procedure KSC-KP-03 (Human resource). Document KSC-KM-02 provides organisation chart. Approved staff are listed in KSC-KP-02-F02.  Organisation chart- See Appendix B.

Training records are documented in KSC-KP-03-F02. Some examples for training in IEC 60079-0, -1, -2, Decision sheets 2022/01, 2021/004 to 006; in ISO 80079-36; ISO 80079-37; IEC 60079-33 were observed

Competence matrix for ExCB and ExTL were provided. Training records and evaluation records were sighted for interviewed staff and found to be satisfactory.

Details of staff competency are included in the site assessment report.

## Complaints and appeals (including appeals to IECEx)

Document procedure KSC-KPA-01 Clause 4.16 addresses complaints and appeals. The procedure address provision of appeals to IECEx and meets the requirements of IECEx scheme.

## Impartiality

Each laboratory staff makes a documented pledge to ensure impartiality which is documented in KSC-KP-01-F01 and confidentiality which is documented in KSC-KP-01-F10 according to procedure KSC-KP-01. Risk assessment including impartiality of laboratory is conducted according to procedure KSC-KP-19 and the assessment result is recorded in KSC-KP-19-01.It is assessed to meet the requirements of IECEx.

## Active involvement in development of Decision Sheets

The ExTL has regular meetings, which are coordinated by both technical manager and quality manager of ExTL, where decision sheets are discussed. This is found to meet the requirements of IECEx.

## Special facts to be noted

None.

## Supporting documentation

Copies of additional supporting information for this assessment have been provided to the applicant and the IECEx Secretariat.  These are included in a site assessment report or provided separately and include:

Details of issues raised and how these have been resolved

Checklist for ISO/IEC 17025

Completed Technical Capability Document (TCD)

Photos of the facilities/tests witnessed are included in the above TCD

Information on competencies

Information on contracting/subcontracting

Assessors’ notes

Other

## Recommendations

Based on the assessment performed on 21 July, 22 July, 26 July, 27 July and 29 July 2022, KSC is recommended for continued acceptance in the IECEx scheme as:

An ExTL in the IECEx Certified Equipment Scheme

This is according to the scope of the standards listed in this document.

Gordana Ostojic

IECEx Lead Assessor

Date:  2022-12 01

# ExCB for IECEx Certified Equipment Scheme – N/A

# ExTL for IECEx Certified Equipment Scheme

## Assessment references

### General references

- IECEx02 IECEx Certified Equipment Scheme covering equipment for use in explosive atmospheres – Rules of Procedure

- IECEx OD003-2 Assessment, surveillance assessment and re-assessment of ExCBs and ExTLs operating in the IECEx 02, IECEx Certified Equipment Scheme

- IECEx OD009 Issuing of CoCs, ExTRs and QARs

- ISO/IEC 17025 General requirements for the competence of testing and calibration laboratories

- IECEx OD 018 Harmonised check list for testing and calibration laboratories ISO/IEC 17025

- IECEx TCD 60079, ISO 80079 Series and ISO 16852 Technical Capability Document

- ExTAG decision sheets (DSs)

- IECEx OD 202 IECEx Certified Equipment Scheme – IECEx Proficiency Testing Program

### Additional references applied for this assessment

- IECEx OD 060 IECEx Guide for Business Continuity – Management of Extraordinary Circumstances or Events Affecting IECEx Certification Schemes and Activities

IECEx OD 280 IECEx Certified Equipment Scheme – Guide to Certification of Non-electrical  Equipment and Protective Systems

IECEx OD233 - IECEx Certified Equipment Scheme - Assessment of Ex “s" Equipment

## ExTL persons interviewed

| Name | Position |
| --- | --- |
| Hyunli Shin | Quality manager |
| Daesik, Moon | Senior Engineer |
| Krzysztof Sinka | Senior Engineer |

## Associated ExCB(s)

KSC POLAND Sp. z o.o.

Eurofins E&E CML Limited (CML)

## Organisation

### Names, titles and experience of the senior executives

| Name | Title | Experience (years) |
| --- | --- | --- |
| Sanghee Kim | CEO | 20 years / 16 years in Ex |
| Hyonggil Kim | Senior Engineer, Technical Manager | 21 years / 12 years in Ex |
| Daesik, Moon | Senior Engineer | 21 years / 21 years in Ex |
| Krzysztof Sinka | Senior Engineer | 17 years / 15 years in Ex |
| Behzad Nejad | Senior Engineer | 17 years / 17 years in Ex |
| Hyunil, Shin | Senior Engineer, Quality Manager | 22 years / 8 years in Ex |

### Name, title and experience of the quality management representative

| Name | Title | Experience (years) |
| --- | --- | --- |
| Hyunil, Shin | Senior Manager | 22 years / 8 years in Ex |

### Other employees in ExTL activity

| Name | Title/responsibility | Experience in Ex (years) |
| --- | --- | --- |
| Hyonggil Kim | Senior Engineer, Technical Manager | 21 years / 12 years in Ex |
| Daesik, Moon | Senior Engineer | 21 years / 21 years in Ex |
| Krzysztof Sinka | Senior Engineer | 17 years / 15 years in Ex |
| Behzad Nejad | Senior Engineer | 17 years / 17 years in Ex |
| Kwangjin Ka | Research Engineer / Testing Engineer | 16 years / 9 years in Ex |
| Hyunil, Shin | Senior Engineer, Quality Manager | 22 years / 8 years in Ex |
| Suna Ryu | Research Engineer / Testing Engineer | 17 years / 8 years in Ex |
| Hoyoung Seo | Assistant Engineer / Testing Engineer | 8 years / 8 years in Ex |
| Jihyun, Son | Assistant Engineer / Testing Engineer | 5 years / 5 years in Ex |
| Yohan, Ko | Assistant Engineer / Testing Engineer | 6 years / 6 years in Ex |
| Seonam Park | Junior Engineer / Testing Engineer | 2 years / 2 years in Ex |
| Eun Hyang Um | Manager / Marketing | 5 years / 1 yeas |

## Organizational structure

Organisational chart is presented in Annex B. Annex A is the Scope of IECExTL Certified equipment Scheme.

## Resources

The ExTL has demonstrated that they have adequate resources, appropriate procedures, working instructions and test facilities for the proposed scope of  Ex protected equipment under IECEx Scheme.

## Test reports issued

Number of test reports (ExTRs) issued under for the preceding two years for each type of protection.

| Standard numbers | Type of protection or other identifying information | Number of issued reports (ExTRs) (for last 2 years) |  | Total |
| --- | --- | --- | --- | --- |
|  |  | 2021 | 2022 |  |
| IEC 60079-1 | flameproof | 28 | 18 | 46 |
| IEC 60079-2 | pressurised | 14 | 3 | 17 |
| IEC 60079-5 | powder | 0 | 0 | 0 |
| IEC 60079-6 | oil | 0 | 0 | 0 |
| IEC 60079-7 | Increased safety | 16 | 9 | 25 |
| IEC 60079-11 | Intrinsic safety | 6 | 8 | 14 |
| IEC 60079-13 | Pressurised rooms | 0 | 0 | 0 |
| IEC 60079-15 | Type ‘n’ | 4 | 3 | 7 |
| IEC 60079-18 | encapsulation | 1 | 0 | 1 |
| IEC 60079-26 | Zone 0 | 1 | 0 | 1 |
| IEC 60079-28 | Optical | 3 | 1 | 4 |
| IEC/IEEE 60079-30-1 | Resistance trace heating | 1 | 0 | 1 |
| IEC 60079-31 | Dust t | 19 | 6 | 25 |
| IEC 60079-36 | Non-electrical | 0 | 0 | 0 |
| IEC 60079-37 | Non-electrical | 0 | 0 | 0 |
| IEC 60079-33 | Special | 1 | 0 | 1 |

## National accreditation

A2LA accreditation certificate is valid until 2024-09-30, certificate number is 4956.02.

Copy of ISO 17025 accreditation is shown in Annex C.

## Calibration

Records of Calibration reports for the instruments used in tests witnessed during audit were reviewed and found satisfactory. ExTL calibration system meets the requirements of ISO/IEC 17025 and IECEx.

## Tests witnessed during the assessment visit

| The following tests were witnessed during this assessment: Standard and edition | Clause number | Test | Comments |
| --- | --- | --- | --- |
| IEC 60079-0 Ed 7 | Clause 26.14 | Measurement of capacitance | Satisfactory |
| IEC 60079-0 Ed 7 | Clause 26.4.2 | Impact tests on enclosure with non-metallic parts | Satisfactory |
| IEC 60079-31 Ed 2 | Clause 6.1.1.4 | IP 6X - Dust test for ta | Satisfactory |
| IEC 60079-0 Ed 7 | Clause 26.4.5 | IPX6 – Water test | Satisfactory |
| IEC 60079-1 Ed 7 | Clause 15.2.2 | Pressure determination for IIC | Satisfactory |
| IEC 60079-1 Ed 7 | Clause 15.2.3 | Over-pressure test | Satisfactory |
| IEC 60079-7 Ed 5 | Clause 6.8 | Temperature-rise test on General purpose connection and junction boxes | Satisfactory |
| IEC 60079-28 Ed 2 | Clause 5.2.2.2 | Measurement of Optical power | Satisfactory |
| IEC 60079-11 Ed 6 | Clause 10.5.3 | Spark ignition and surface temperature on battery | Satisfactory |
| IEC 60079-15 Ed | Clause 11.3 | Restricted-breathing testing | Satisfactory |

## Participation in IECEx Proficiency Testing Programs

| Year(s) of participation | IECEx Proficiency Testing program | General information about results |
| --- | --- | --- |
| 2019-2020 | Program 9 “Battery Testing” | Satisfactory. |
| 2019-2020 | Program 10 “Tests of Enclosures” | Satisfactory. |
| 2021-2022 | PTB_Flameproof Joints Test Round 2021 | Satisfactory. |
| 2021-2022 | PTB_Small component Test Round 2021 | Under proceeding to Phase II |

## Comments (including issues found during assessment)

KSC has the necessary resources and quality system in place for their scope as an IEC ExTL.

There were some findings related to test records details and the set-up verification steps during testing. All issues were addressed to the satisfaction of the IECEx Assessor and now meet the requirements of the IECEx.

5    ATF for IECEx Certified Equipment Scheme – N/A

# ExCB for Certified Service Facilities Scheme – N/A

# IECEx Conformity Mark Licensing Scheme – N/A

# Annexes

Scope for IECEx Certified Equipment Scheme

Current standards

| Number | Title | Comments |
| --- | --- | --- |
| IEC 60079-0 Edition 7.0 | Explosive atmospheres - Part 0: Equipment - General requirements | Already in scope |
| IEC 60079-1 Edition 7.0 | Explosive atmospheres - Part 1: Equipment protection by flameproof enclosures “d” | Already in scope |
| IEC 60079-2 Edition 6.0 | Explosive atmospheres - Part 2: Equipment protection by pressurized enclosure “p’ | Already in scope |
| IEC 60079-5 Edition 4.1 | Explosive atmospheres - Part 5: Equipment protection by powder filling “q” | Already in scope |
| IEC 60079-6 Edition 4.1 | Explosive atmospheres - Part 6: Equipment protection by oil immersion “o” | Already in scope |
| IEC 60079-7 Edition 5.1 | Explosive atmospheres - Part 7: Equipment protection by increased safety "e" | Already in scope |
| IEC 60079-11 Edition 6.0 | Explosive atmospheres - Part 11: Equipment protection by intrinsic safety “i” | Already in scope |
| IEC 60079-13 Edition 2.0 | Explosive atmospheres - Part 13: Equipment protection by pressurized room "p" and artificially ventilated room "v" | Already in scope |
| IEC 60079-15 Edition 5.0 | Explosive atmospheres – Part 15: Equipment protection by type of protection "n" | Already in scope |
| IEC 60079-18 Edition 4.1 | Explosive atmospheres – Part 18: Equipment protection by encapsulation “m” | Already in scope |
| IEC 60079-26 Edition 3.0 | Explosive atmospheres - Part 26: Equipment with equipment protection level (EPL) Ga | Already in scope |
| IEC 60079-28 Edition 2.0 | Explosive atmospheres - Part 28: Protection of equipment and transmission systems using optical radiation | Already in scope |
| IEC/IEEE 60079-30-1 Edition 1.0 | Explosive atmospheres – Part 30-1: Electrical resistance trace heating – General and testing requirements | Already in scope |
| IEC 60079-31 Edition 3.0 | Explosive atmospheres – Part 31: Equipment dust ignition protection by enclosure "t" | Already in scope |
| IEC 60079-33 Edition 1.0 | Explosive atmospheres – Part 33: Equipment protection by special protection “s” | Already in scope |
| IS0 80079-36 Edition 1.0 | Explosive atmospheres - Part 36: Non-electrical equipment for explosive atmospheres – Basic method and requirements | Already in scope |
| ISO 80079-37 Edition 1.0 | Explosive atmospheres - Part 37: Non-electrical equipment for explosive atmospheres – Non electrical type of protection constructional safety ”c” control of ignition source ”b”, liquid immersion ”k” | Already in scope |
| IEC TS 60079-46 Edition 1.0 | Explosive atmospheres – Part 46 - Equipment assemblies | Already in scope |

Superseded standards

| IEC 60079-31 Edition 2.0 | Explosive atmospheres – Part 31: Equipment dust ignition protection by enclosure "t" | Already in scope |
| --- | --- | --- |
| IEC 60079-26 Edition 2.0 | Explosive atmospheres - Part 26: Equipment with equipment protection level (EPL) Ga | Already in scope |

Overall Organisation Chart

Accreditation Certificate for ISO/IEC 17025